PRODUCT NAME:

Femobiome II Multiplex qPCR Kit

INTENDED USE:

Femobiome II Multiplex qPCR Kit is an in vitro Nucleic Acid Test (NAT). Femobiome II Multiplex qPCR Kit is designed for quantitative detection of DNA of normobiota representatives, qualitative and quantitative detection of DNA of pathogenic microorganisms  in biomaterial from female reproductive tract (epithelial swab from the mucous membrane of cervical canal, vagina) by real-time PCR.

Headquarters: Neighborhood of Dumlupınar, Nur St. No: 7/3 Post Code: 41250 Kartepe/Kocaeli/TURKEY

Branch: TUBITAK MAM Technology Free Zone, Neighborhood of Barıs SB, 5002 Street, New Tech. Building, Blok A, No: 4A/ Z09 Post Code: 41400

Gebze/Kocaeli/TURKEY

Trd Register Number: 34218

Email: info@genomer.com.tr

https://genomer.com.tr

Purpose

This document constitutes the Instructions for Use (IFU) for the Femobiome II Multiplex qPCR Kit, an In Vitro Diagnostic (IVD) medical device.

The European Union's In Vitro Diagnostic Medical Devices Regulation (IVDR) 2017/746 aims to ensure a high level of health and safety protection for patients and users throughout the entire lifecycle of the device.

These Instructions for Use have been prepared to fully satisfy all obligations stipulated under Annex I: General Safety and Performance Requirements of the IVDR, specifically within Chapter III: Information Supplied by the Manufacturer (Article 20).

The primary objective is to ensure the safe, accurate, and effective use of the device for its intended purpose. These instructions encompass all necessary technical, procedural, and legal information for laboratory professionals, including device performance, limitations, potential risks, and appropriate sample management.

Regulatory Compliance and Transparency

This guide not only employs pictograms and symbols strictly in accordance with national and international standards (specifically IVDR, EN ISO 15223-1, and EN ISO 18113) but also transparently presents all legally required device identifiers (UDI, Manufacturer, Authorized Representative).

Warning: It is imperative to read, understand, and meticulously adhere to these instructions in their entirety prior to using the device to ensure safe and accurate results. Failure to comply may lead to erroneous results, potential hazards, or invalidation of the CE marking.

| I. DESCRIPTIVE INFORMATION |  |  |  |
| --- | --- | --- | --- |
| 1. | Device Name, Trade Name, and Model | Article 20.3(a), 20.4.1(b) | Femobiome II Multiplex qPCR Kit |
| 2. | In Vitro Diagnostic Device | Article 20.3(b) |  |
| 3. | Manufacturer Information | Article 20.3(d) | GENOMER BIOTECHNOLOGY R&D IND. TRD. LIMITED COMPANY, Headquarters: Neighborhood of Dumlupınar, Nur St. No: 7/3 Post Code: 41250 Kartepe/Kocaeli/TURKEY Branch: TUBITAK MAM Technology Free Zone, Neighborhood of Barıs SB, 5002 Street, New Tech. Building, Blok A, No: 4A/ Z09 Post Code: 41400 Gebze/Kocaeli/TURKEY Trd Register Number: 34218 Email: info@genomer.com.tr Web:https://genomer.com.tr |
| 4. | Authorized Representative Information | Article 20.3(e) | Required only if the manufacturer is located outside the EU. |
| 5. | Lot/Batch Number | Article 20.3(f) | UL- FBII-XX |
| 6. | Expiry/Use-by Date | Article 20.3(g) | YYYY/AA/GG |
| 7. | Catalog Number | Article 20.3(c) | M-FBII-0701v01-12 |
| 8. | Unique Device Identifier (UDI) Reference | Article 20.3(n) | 868512670MFBIIKF |

The results obtained in accordance with the procedures and analytical acceptance criteria specified in the aforementioned Instructions for Use (IFU) guide confirm the conformity of Femobiome II Multiplex qPCR Kit device with the General Safety and Performance Requirements (Annex I).

| II. INTENDED USE, SCOPE, AND PRINCIPLE |  |  |
| --- | --- | --- |
| 9. | Article 20.4.1(c) | Intended Use and Function |

Femobiome II Multiplex qPCR Kit is an in vitro Nucleic Acid Test (NAT). Femobiome II Multiplex qPCR Kit is designed for quantitative detection of DNA of normobiota representatives (Lactobacillus spp., Lactobacillus iners, Lactobacillus crispatus, Lactobacillus jensenii/Lactobacillus mulieris, Lactobacillus gasseri/Lactobacillus paragasseri, Bifidobacterium spp.), opportunistic microorganisms (Ureaplasma urealyticum, Ureaplasma parvum, Mycoplasma hominis, Candida spp., Candida albicans, Staphylococcus spp., Streptococcus spp., Streptococcus agalactiae, Enterobacteriaceae, Enterococcus spp., Haemophilus spp., Gardnerella vaginalis, Fannyhessea vaginae, Mobiluncus spp., Anaerococcus spp., Peptostreptococcus spp., Bacteroides spp./Porphyromonas spp./Prevotella spp., Sneathia spp./Leptotrichia spp./Fusobacterium spp., Megasphaera spp./Veilonella spp./Dialister spp., BVAB1/BVAB2/BVAB3), qualitative and quantitative detection of DNA of pathogenic microorganisms (Chlamydia trachomatis, Neisseria gonorrhoeae, Mycoplasma genitalium, Trichomonas vaginalis, HSV-1, HSV-2, CMV,  HPV 16,  HPV 18, HPV 45, HPV 31/33/35/39/51/52/56/58/59/66/68) in biomaterial from female reproductive tract (epithelial swab from the mucous membrane of cervical canal, vagina) by real-time PCR.

Population and demographic aspects: the reagent kit is designed for examination of post-pubertal women.

Potential users: qualified personnel trained in molecular diagnostic methods. Apply the kit only as directed in the present instruction for use.

| II. INTENDED USE, SCOPE, AND PRINCIPLE |  |  |
| --- | --- | --- |
| 10. | Article 20.4.1(e) | Test Principle - Method |

Method: multiplex polymerase chain reaction (PCR) with results detection in real-time; quantitative and qualitative multiplex analysis.

The amplification process consists of a series of repeated cycles of temperature denaturation of DNA, annealing of primers with complementary sequences, and subsequent elongation of polynucleotide chains with Taq polymerase.

To increase the sensitivity and specificity of the amplification reaction, hot start is used. For package S, hot start is provided by reaction mixture preparation consisting of two layers separated by a layer of paraffin. The PCR starts only when paraffin is melted. “Hot” start for packages A and A-TL is provided by blocking the activity of Taq polymerase with antibodies. The enzyme activates only after preheating the reaction mixture at 94°C. This prevents the nonspecific annealing of primers on the DNA target at low temperatures during the initial heating of the test tube.

DNA probes, each containing a fluorescent label and a fluorescence quencher, are introduced into the PCR mix. When a specific product is formed, the DNA probe is destroyed and the quencher stops affecting the fluorescent label, which leads to an increase in the fluorescence level. The number of hydrolyzed probes (and hence the fluorescence level) increases in proportion to the number of specific amplicons formed. The fluorescence level is measured at each amplification cycle in real time.

The reagent kit includes PCR mixes specific for lactobacilli, pathogenic and opportunistic microorganisms, DNA of all bacteria (Total Bacterial Load - TBL) and human genomic DNA (HGD).

PCR mixes (except for tube No. 3 of strip No. 1 and tube No. 4 of strip No. 2) include internal control (IC) intended for quality assessment of PCR in each tube.

A marker oligonucleotide with a fluorescent dye Rox is added into tube No.2 of strip No.1 and into tube No.3 of strip No.2 to control the position of strips in the thermoblock. After amplification, the software compares the marker position specified by the operator with its real position, and warns the operator in case of incorrect strip positioning. In this case the operator needs to verify the position of the strips in the thermoblock (the first tube of Strip No. 1 and the second tube of Strip No. 2 are marked with a blue buffer) and correct the tube IDs in the protocol.

The DNA probes used for detection of specific analytes include fluorescent tags: Fam, Rox, Hex, Cy5, Cy5.5. The DNA probe used for detection of the amplification product of internal control includes fluorescent dye Hex.

The use of several fluorescent dyes allows to reduce the number of tubes and biomaterial required for the assay by simultaneously registering the results of multiple amplification reactions.

Table 1 shows the detected microorganisms, color coding, and detection channels for amplification products.

| Tube No. |  |  |  |  |  | PCR mix color coding |
| --- | --- | --- | --- | --- | --- | --- |
|  | Fam | Hex | Rox | Cy5 | Cy5.5 |  |
| 1 | Total bacterial load (TBL) | IC | - | - | Gardnerella vaginalis | Blue |
| 2 | Lactobacillus spp. | IC | Marker | - | - | Colorless or pink |
| 3 | Lactobacillus crispatus | Lactobacillus iners | Lactobacillus jensenii/ Lactobacillus mulieris | Lactobacillus gasseri/ Lactobacillus paragasseri | - |  |
| 4 | Staphylococcus spp. | IC | Streptococcus spp. | Streptococcus agalactiae | - |  |
| 5 | Mobiluncus spp. | IC | - | Haemophilus spp. | - |  |
| 6 | Anaerococcus spp. | IC | Fannyhessea vaginae | - | - |  |
| 7 | Bacteroides spp./ Porphyromonas spp./ Prevotella spp. | IC | Enterobacteriaceae | Enterococcus spp. | - |  |
| 8 | Sneathia spp./ Leptotrichia spp./ Fusobacterium spp. | IC | - | - | - |  |
| 1 | Megasphaera spp./ Veilonella spp./Dialister spp. | IC | - | - | - | Colorless or pink |
| 2 | BVAB1/BVAB2/BVAB3 | IC | - | - | - | Blue |
| 3 | Peptostreptococcus spp. | IC | Marker | Bifidobacterium spp. | - | Colorless or pink |
| 4 | Mycoplasma hominis | Ureaplasma urealyticum | Ureaplasma parvum | Human genomic DNA | - |  |
| 5 | Candida spp. | IC | Candida albicans | - | - |  |
| 6 | Trichomonas vaginalis | IC | Mycoplasma genitalium | Chlamydia trachomatis | Neisseria gonorrhoeae |  |
| 7 | HSV-2 | IC | CMV | HSV-1 | - |  |
| 8 | HPV 16 | IC | HPV 31/33/35/39/51/ 52/56/58/59/66/68 | HPV 18 | HPV 45 |  |

The automatic analysis is available on G-Quant Real Time PCR ThermoCylers (see  https://genomer.com.tr/  for  available   supply   options). The current software version is available for download at https://genomer.com.tr/.

The indicator cycle number and PCR-efficiency are used to calculate the amount of target analyte. To determine the indicator cycle, the crossing point method (Ср) is used. The Cp method is based on the calculation of the maximum of the second derivative of a fluorescent graph. PCR efficiency is an indicator of the increase in the number of DNA fragments within one amplification cycle.

The decimal logarithm (Lg) of initial DNA concentration in a sample (number of DNA target copies per

1.0 mL of DNA sample) is calculated using indicator cycle comparison method, also known as ΔΔCt method. The ΔΔCt method makes it possible to determine the initial amount of target DNA in a sample without the need to construct a calibration curve.

Calculation is carried out according to the formula:

𝐿𝑜𝑔(𝑁𝑂𝑎) = 𝐿𝑜𝑔(𝐸) × (𝐶𝑝 − 𝐶𝑝𝑎),	(1)

where N0a is the initial number of target DNA copies in 1.0 mL of DNA sample;

E – PCR efficiency;

Cp – indicator cycle for the sample containing a single DNA molecule

Cpa – indicator cycle for the test sample.

The assay consists of the following steps: DNA extraction (DNA sample preparation) and DNA PCR amplification with detection of the results in real time using Femobiome II Multiplex qPCR Kit.

| II. INTENDED USE, SCOPE, AND PRINCIPLE |  |  |
| --- | --- | --- |
| 11. | Article 20.3(k) | Adequate Number of Tests |

Femobiome II Multiplex qPCR Kit is a package containing the following:

| Reagent | Description | Total volume | Amount |
| --- | --- | --- | --- |
| Paraffin-sealed PCR mix. Strip No. 1 | Colorless, pink or blue transparent liquid under waxy white fraction | 1.92 mL (20 μL in each) | tubes, 12 strips of 8 |
| Paraffin-sealed PCR mix. Strip No. 2 | Colorless, pink or blue transparent liquid under waxy white fraction | 1.92 mL (20 μL in each) | tubes, 12 strips of 8 |
| Taq polymerase solution | Colorless transparent liquid | 2.0 mL (1.0 mL in each) | 2 tubes |
| Mineral oil | Colorless transparent viscous oily liquid | 4.0 mL (1.0 mL in each) | 4 tubes |
| Positive control | Colorless transparent liquid | 320 μL | 1 tube |
| Strip caps | 24 strips of 8 |  |  |

All components are ready for use.

Femobiome II Multiplex qPCR Kit is designed for 12 tests (no more than 4 runs), which includes analysis of test samples, negative and positive controls.

ADDITIONAL REAGENTS AND EQUIPMENT REQUIRED

- UV PCR cabinet

- Real-time PCR thermal cycler

- Vortex mixer

- Vortex rotor for 0.2 mL strip tubes

- Refrigerator or a cooling chamber

- Freezer

- Tube rack for 0.2 mL strip tubes

- Tube rack for 1.5 mL tubes

- Single channel pipettes (dispensers covering 2.0-20, 20-200, 200-

- 1,000 μL volume range);

- RNase- and DNase-free aerosol barrier pipette tips (volumes 10

- μL, 20 μL, 200 μL, 1000 μL)

- Pipette rack

- RNase- and DNase-free 1.5 mL microfuge tubes with caps

- Powder-free surgical gloves

- Container for used pipette tips, tubes and other consumables

- STOR F Transport Medium

- Transport media for liquid cytology: PreservCyt®, Hologic Inc., USA; BD SurePath™ Liquid-Based Pap Test, Becton, Dickinson and Company, USA; EASYPREP, YD Diagnostics, South Korea; CellPrep, CP Biodyne, South Korea; Сell Preservative Solution, Human Lituo Biotechnology Co.Ltd, China

- Medical swabs, types А, А1

- Self-sampling devices: Qvintip self-sampling device, Aprovix AB, Sweden; Copan swab system, Copan Italia S.p.A., Italy

- NA extraction kit: PREP-NA PLUS DNA/RNA Extraction Kit and PREP-MB MAX DNA Extraction Kit

- Software: The most recent version of the G-Quant Real Time PCR ThermoCylers software can be downloaded from https://genomer.com.tr/.

| III. SAFETY AND PRECAUTIONS |  |  |
| --- | --- | --- |
| 12. | Article 20.4.1(n) | Warnings and Precautions |
| 13. | Article 20.4.1(p) | Biohazard Warning |
| 14. | Article 20.4.1(o) | Hazardous Chemical Substances |
| 15. | Article 20.1(f) | Restriction on Use |

- For in vitro diagnostic use only.

- Molecular biology procedures, such as nucleic acid extraction, PCR-amplification and detection require qualified staff to avoid the risk of erroneous or unreliable results, especially due to the degradation of nucleic acids contained in the samples or sample contamination by amplification products.

- Wear powder-free single-use surgical gloves. Wear work clothes and personal protective equipment while working with pathogenic microorganisms. The work clothes and personal protective equipment must be suitable for work to be performed and comply with health and safety requirements.

- Avoid any direct contact with the biological samples, reagents and materials used to carry out the assay. Avoid producing spills or generating aerosols. Do not eat/drink components of the kit. Do not inhale gas/fumes/vapor/aerosols produced by the components of the kit. Avoid contact with eyes.

- Samples must be handled under a laminar flow hood. Tubes containing different samples must never be opened at the same time.

- Pipettes used to handle samples must only be used for one purpose. The pipettes must be of positive displacement type or be used with aerosol barrier pipette tips. The tips employed must be sterile, free from DNases and RNases and free from DNA and RNA.

- The reagents must be handled under a laminar flow hood. The reagents required for amplification must be prepared in such a way to be utilized in a single session.

- Handle and dispose of all biological samples, reagents and materials used to carry out the assay as if infectious1, 2. Any material being exposed to biological samples must be treated with disinfecting solution for at least 30 minutes or autoclaved for 1 hour at 121°C before disposal.

- All of the liquid solutions are designed for single use and cannot be used more than once in amplification reactions.

- Only use the reagents provided in the kit and those recommended by the manufacturer. Do not mix reagents from different batches. Do not use reagents from third party manufacturers’ kits.

- All laboratory equipment and tools, including pipettes, test tube racks, laboratory glassware, lab coats, bouffant caps, gloves, etc., as well as reagents must be strictly stationary. It is not allowed to move them from one room to another. Equip separate areas for the extraction/preparation of amplification reactions and for the amplification/detection of amplification products. Never transfer lab coats, gloves and tools from the area designed for amplification/detection of the amplification products to the area designed for extraction/preparation of amplification reactions. Never introduce amplification products in the area designed for extraction/preparation of amplification reactions.

- Do not open the tubes after amplification. Avoid producing accidental spills of the amplification products. Dispose of all PCR waste materials (tubes, tips etc.) only in a closed form in a specialized sealed container with disinfectant solution. Waste materials must be removed in accordance with laboratory internal procedures, and with national and international standards.

- Working surfaces, as well as rooms where NA extraction and PCR are performed, must be disinfected with bactericidal irradiators (UVGI) for 30 minutes before and after the assay. All surfaces in the laboratory (test tube racks, equipment, tools, etc.) must be treated with disinfecting solution daily.

Emergency actions

Eye Contact: If any component of the kit enters the eyes, flush the eyes gently using potable running water for 15 minutes or longer, making sure that the eyelids are held open. If pain or irritation occurs, seek medical attention.

Skin Contact: If any component of this kit comes into contact with the skin and causes discomfort, remove any contaminated clothing. Rinse the affected area with plenty of soap and water. If pain or irritation occurs, seek medical attention.

Ingestion: If any component of this kit is ingested, rinse the mouth with plenty of potable water. If irritation or discomfort occurs, seek medical attention.

Do not use the kit:

−	If the transportation and storage conditions have been violated;

−	If the appearance of the reagents does not correspond to the product documentation;

−	If the packaging of the kit components is breached;

−	After the expiry date.

Adverse health effects are NOT anticipated from routine use of this kit in compliance with the current instruction for use.

| III. SAFETY AND PRECAUTIONS |  |  |
| --- | --- | --- |
| 16. | Article 20.4.1(j) | Special Storage Conditions |

Expiry date – 12 months from the date of manufacture.

Storage conditions

- All components of Femobiome II Multiplex qPCR Kit must be stored at temperatures from 2°C to 8°C over the storage period.

- Paraffin-sealed PCR mix must be stored away from light over the storage period.

Transport conditions

- Transportation of the reagent kit is carried out in thermoboxes with ice packs by all types of roofed transport at the temperature inside the container corresponding to the storage conditions of the kit components.

- It is allowed to transport the kit in thermoboxes with ice packs by all types of roofed transport at the temperature inside the thermoboxes from 2°С to 25°С for no longer than 5 days.

Shelf-life of the kit following the first opening of the primary container

- All components of the kit must be stored in a refrigerator or a cooling chamber at temperatures from 2°C to 8°C over the storage period;

- Paraffin sealed PCR-mix must be stored in a refrigerator or a cooling chamber at temperatures from 2°C to 8°C and out of light over the storage period;

WARNING! Kits stored with violation of storage conditions must not be used.

An expired Femobiome II Multiplex qPCR Kit must not be used.

We strongly recommend following the current instructions for use in order to obtain accurate and reliable results.

| IV. TEST PROCEDURE AND QUALITY CONTROL |  |  |
| --- | --- | --- |
| 17. | Article 20.4.1(d) | Sample Type |

Femobiome II Multiplex qPCR Kit is designed to detect DNA extracted from epithelial swabs from mucous membrane of the cervical canal and vagina.

General requirements

PCR analysis is a direct diagnostic method, therefore, the collection of biological material must be carried out from the site of infection localization. The biomaterial type is selected by the physician based on the collected anamnesis and clinical picture of the disease.

The quality of sample collection, sample storage, transport and pretreatment is crucial for obtaining correct results. Incorrect sample collection may lead to unreliable results and, therefore, to the need to repeat the sample collection procedure.

If biomaterial from several biotopes is required, repeat the procedure using a new swab for each biotope. Use RNase- and DNase-free aerosol barrier pipette during biomaterial preparation and NA extraction.

To prevent contamination, only open one tube cap at a time and close it before proceeding to the next tube.

Sample collection and preparation

WARNING! Sample pretreatment may be required before DNA extraction.

Sample collection is performed using specialized medical devices in accordance with the procedure outlined below. Additionally, consult the instructions for use for NA extraction reagent kit and self- sampling device before collecting the sample.

| BEFORE THE TEST, IT IS NOT ALLOWED TO | IF THE GUIDELINES ARE NOT FOLLOWED, biomaterial can be collected no sooner than |
| --- | --- |
| Perform syringing and vaginal douching with antiseptics | 24 hours |
| Use tampons |  |
| Use PCR-inhibiting substances, such as ultrasound gel, heparin, chlorhexidine and other chlorine- containing medicinal products |  |
| Undergo transvaginal ultrasound examination |  |
| Engage in protected sexual contacts |  |
| Undergo colposcopy | 48 hours |
| Engage in unprotected sexual contacts | 72 hours |
| Take and/or use antibacterial medications | 2 weeks FEMOFLOR II can be prescribed to selected patients on antibiotic therapy, for instance, for therapy monitoring. However, the test results should be interpreted with care by the doctor based on the patient’s medical history, as antibiotics can influence the outcome of the test. |
| Take and/or use probiotics |  |

Table 6: Preparation for biomaterial collection

WARNING! Before obtaining an epithelial cell swab from the posterior vaginal vault and cervical canal, the free-flowing secretion should be removed with a sterile cotton swab.

Vaginal sampling

The sample must be taken prior to physical examination. Speculum can be treated with warm water before the procedure. Antiseptics must not be used for speculum treatment. The sample must be taken from the lateral or posterolateral vaginal wall.

For virginal women, the sample must be taken from the vestibular mucous membrane, or, in special cases, from the posterior vaginal wall through the hymenal rings.

Cervical sampling

Remove mucus with a cotton swab prior to sampling and treat the cervix with sterile physiological saline solution. Carefully insert a sampling swab into the cervix to a depth of 0.5-1.5 cm. Avoid contact with the vaginal wall when removing the swab.

Transport and storage of the samples

Transport and storage conditions for epithelial swabs from the mucous membrane of the cervical canal and vagina are outlined in the instructions for use to the recommended NA extraction reagent kits or transport media used for transport and storage of samples.

Biomaterial preparation for DNA extraction

Preparation of epithelial swabs from the mucous membrane of the cervical canal, vagina is carried out in accordance with the instructions for use for the NA extraction reagent kits and for transport media.

For samples taken into BD SurePath™ Liquid-Based Pap Test and CellPrep media, pretreatment is required. It is recommended to use NA Extraction Biomaterial Pretreatment Reagent Kit (PREP-PK) manufactured by DNA-Technology Research & Production, LLC, according to instructions for use.

For samples taken into PreservCyt, EASYPREP and Сell Preservative Solution media:

- Thoroughly mix the contents of the vial with the sample in the transport medium (solution) by shaking the vial intensively.

- Transfer 1.0 mL of material into a disposable 1.5 mL plastic tube. Close the tube tightly.

- Centrifuge the tube at RCF(g) 12,000-16,000 for 3 minutes.

- Carefully remove as much supernatant as possible without disturbing the pellet.

- Add STOR-F Transport Medium or sterile saline solution to the precipitate in the volume specified in the instructions for the DNA extraction reagent kit and resuspend the precipitate by pipetting.

When using self-sampling devices, the probe is transferred into a tube with transport medium intended for the transport and storage of biomaterial samples for PCR assay. Further preparation is carried out in accordance with the instructions for use for NA extraction reagent kits or the instructions for use for the transport medium used.

| IV. TEST PROCEDURE AND QUALITY CONTROL |  |  |
| --- | --- | --- |
| 18. | Article 20.3(j) | Refer to Instructions for Use |
| 19. | Article 20.4.1(l) | Procedure Steps |

For DNA extraction from epithelial swabs from mucous membranes of cervical canal and vagina it is recommended to use the authorized NA extraction reagents kits compatible with the biomaterial, for example PREP-NA PLUS DNA/RNA Extraction Kit and PREP-MB MAX DNA Extraction Kit.

The minimal eluate volume for the reagent kits mentioned above is 300 μL.

DNA extraction is performed in accordance with the instructions for use of the corresponding reagent kit; for PREP-NA PLUS DNA/RNA Extraction Kit short method of extraction can be used.

WARNING! Regardless of the DNA extraction kit used, the negative control must go through all stages of DNA extraction. Physiological saline solution can also be used as a negative control in volumes specified in the instruction for use for the corresponding DNA extraction kit.

Test Assay

WARNING!

- The reagents and tubes must be kept away from direct sunlight.

- Closely monitor the completeness of the strips and caps for them. Do not use strip caps from other kits!

1. Mark two strips (Strip No. 1 and Strip No. 2) with paraffin-sealed PCR mix for each test sample, positive (C+) and negative (C-) controls. Note. Two strips are used for one sample.

Example: to test 4 samples, mark 4 Strips No. 1 and 4 Strips No. 2 for test samples; one Strip No. 1 and one Strip No. 2 for “C-“ and one Strip No. 1 and one Strip No. 2 for “C+“. The total number of strips is 12.

2. Shake the tube with Taq polymerase solution on vortex for 3-5 seconds and centrifuge at RCF(g) 1,000-3,000 for 1-3 seconds.

3. Add 10 μL of Taq polymerase solution into each tube. Avoid paraffin layer break.

4. Add one drop (~20 μL) of mineral oil into each tube of the strip. Cover the strips loosely with caps.

5. Shake the tubes with samples, “C+” and “C-” for 3-5 seconds and centrifuge at RCF(g) 1,000-3,000 for 1-3 seconds.

WARNING!

- Before adding DNA and negative controls to the test tubes with the reaction mixture follow the DNA preparation guidelines outlined in the instruction for use for the NA extraction kit.

- For PREP-NA PLUS DNA/RNA Extraction Kit: if the precipitate with the extracted DNA was transferred to new tubes, shake the tubes with DNA sample and negative control on vortex for 3-5 seconds, then centrifuge at RCF(g) 1,000-3,000 for 1-3 seconds.

- For PREP-MB MAX DNA Extraction Kit, centrifuge the tubes with the DNA sample and negative control at RCF(g) 1,000-3,000 for 1-3 seconds (do not shake!), then place the tubes into magnetic rack. If after extraction the precipitate with the extracted DNA was transferred to new tubes, shake the tubes with the DNA sample and negative control on vortex for 3-5 seconds, then centrifuge at RCF(g) 1,000-3,000 for 1-3 seconds.

- Before introducing DNA, only open the caps of the strips that are currently in use and close them before proceeding to the next one to prevent contamination. Close the strips tightly. Use aerosol barrier pipette tips.

6. Add 5.0 μL of DNA sample into each tube of the corresponding labeled strips. Do not add DNA

into the “C+”, “C-” strips. Avoid paraffin layer break. Close the strips tightly.

7. Add 5.0 μL of the negative control (C-) (the one that underwent the whole DNA extraction procedure) into each tube of the corresponding labeled strips. Close the strip tightly.

8. Add 5.0 μL of positive control (C+) into each tube of the corresponding labeled strips. Avoid paraffin layer break. Close the strip tightly.

9. Centrifuge all strips at RCF(g) 1,000-3,000 for 1-3 seconds.

10. Place all strips into the real-time PCR thermal cycler.

11. Launch the operating software for G-Quant Real Time PCR ThermoCylers. Add the corresponding test, specify the number and IDs of the samples, positive and negative controls. Specify the position of the strips in the thermoblock and run PCR. Before the run, compare your selected program to the values provided in Table 7.

| Step | Temperature, °C | Min | Sec | Number of cycles | Optical measurement | Step type |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | 80 | 0 | 30 | 1 |  | Cycle |
|  | 94 | 1 | 30 |  |  |  |
| 2 | 94 | 0 | 30 | 5 |  | Cycle |
|  | 64 | 0 | 15 |  | √ |  |
| 3 | 94 | 0 | 10 | 45 |  | Cycle |
|  | 64 | 0 | 15 |  | √ |  |
| 4 | 94 | 0 | 5 | 1 |  | Cycle |
| 5 | 25* | … | … | Temperature hold |  | Temperature hold |
| *temperature hold at 10°С is allowed |  |  |  |  |  |  |

Table 1: PCR program for G-Quant Real Time PCR ThermoCylers

| IV. TEST PROCEDURE AND QUALITY CONTROL |  |  |
| --- | --- | --- |
| 20. | Article 20.4.1(m) | Quality Control (QC) Procedure |

- Femobiome II Multiplex qPCR Kit contains a positive control which is produced using genetic engineering techniques and contains the target DNA sequence.

- To reveal possible contamination, a negative control is required.

WARNING! The negative control must go through all stages of DNA extraction

- The kit contains a PCR mix for human genomic DNA (HGD). HGD estimates the amount of human genomic DNA in the tube.

- The PCR mix contains control indicator (tube 4 of Strip No. 2, Cy5 detection channel) which is based on the amount of human genomic DNA (HGD) present in the tube.

- The PCR mix also contains Total Bacterial Load (TBL) in tube 1 of Strip No. 1, Fam channel, which quantifies the total amount of microorganisms present in the tube.

- The PCR mix from the kit contains an internal control (IC) in tubes 1-2, 4-8 of Strip No. 1 and tubes 1-3, 5-8 of Strip No. 2, Hex detection channel. IC is an artificial plasmid used to assess the quality of PCR performance. IC is absent in tube 3 of Strip No. 1 and tube 4 of Strip No. 2.

The test result is considered invalid when the exponential growth of the fluorescence level for the specific amplification product and for internal control (IC) are not observed.

HGD value must be considered when analyzing results: HGD values lower than 3.5 indicate insufficient amount of sample; the sampling procedure must be repeated.

The test result is considered valid if:

For positive control the test result must be displayed as a decimal logarithm of DNA concentration or as “+”, meaning “detected”. In case of a negative result (“-”, “not detected”) for the positive control the results of the current run are considered invalid. In this case it is required to repeat amplification of all DNA samples.

For negative control the test result must be displayed as “-”, meaning “not detected” for specific amplification target and as “+”, or “detected” for internal control. Otherwise, the results of the current run are considered invalid; decontamination procedure is required.

| V. RESULTS AND PERFORMANCE |  |  |
| --- | --- | --- |
| 21. | Article 20.4.1(m) | Interpretation of Results |

1. Amplification results are calculated automatically by the software that is provided with the thermal cycler.

2. After amplification, the software compares the position of the tubes specified by the operator with the actual position of the marker. Software warns the operator if there is a mismatch. If so, it is necessary to check the position of the strips in the thermoblock (tube 1 of Strip No. 1 and tube 2 of Strip No. 2 are marked with a blue buffer solution) and correct the tube IDs in the protocol.

3. When analyzing the results, the values of TBL, HGD and internal controls must be considered;

3.1. If negative results are obtained in the same tube simultaneously for IC and for target analytes on the corresponding detection channels the assay, the result for that sample is considered invalid (except for tubes without IC). An invalid result may be due to the presence of inhibitors in the DNA sample obtained from biological material; incorrect execution of the assay protocol, inadequate amplification temperature, etc. In this case, it is necessary to either repeat PCR with the same DNA sample, repeat DNA extraction and amplification for this sample, or re-sample the biological material from the patient (performed sequentially until the error is eliminated).

3.2. TBL and HGD values are used to determine whether the amount of biomaterial is sufficient to estimate the relative quantity of normobiota representatives and opportunistic microorganisms. For pathogenic microorganisms the result is considered valid regardless of the TBL and HGD values (10.7).

4. Test results for biomaterial samples containing DNA of target analytes in concentrations above the detection threshold but below the limit of quantification are displayed as “<4.0”.

5. Test results for biomaterial samples containing DNA of target analytes, with the exception of pathogenic microorganisms, will be negative for the respective analytes (“-”) and positive for IC (except for tubes without IC) if target DNA concentrations are below the detection threshold.

6. Results of the run are valid if the following conditions for controls are observed:

6.1. For negative control, a negative result (“-”) must be obtained for the target analytes on the corresponding detection channels and a positive result must be displayed for IC (except for tubes without IC) on the Hex detection channel.

Results below detection threshold for all analytes, except for Chlamydia trachomatis, Neisseria gonorrhoeae, Mycoplasma genitalium, Trichomonas vaginalis, HSV-1, HSV-2,  CMV,  HPV 16,  HPV 18, HPV 45, HPV 31/33/35/39/51 /52/56/58/59/66/68 are displayed as negative (“-“).

If positive result is obtained for the negative control for target analytes the results of the entire run are considered invalid. In this case, measures should be taken to identify and eliminate possible contamination.

6.2. A positive result (5.0±1.0 Lg GE/mL) must be obtained for the positive control. The result obtained for the IC is considered irrelevant in this case.

If a negative result is obtained for a positive control, the results of the entire run are considered invalid. In this case it is necessary to repeat the amplification procedure for all samples.

7. Interpretation of the result consists of the following steps:

- Check the values for TBL and HGD. Lg GE/mL values for TBL and HGD <3.5 for cervical swabs and values for TBL <4.0 and HGD <3.5 for vaginal swabs should be interpreted as an insufficient amount of biomaterial for this assay;

- For normobiota representatives, aerobes and anaerobes the result can be either:

  - negative (“Not detected”; “-”);

  - positive, as “<3.0” Lg GE/mL, if the DNA concentration is below the limit of quantification (LoQ);

  - positive, as “X” Lg GE/mL, if the DNA concentration is above the limit of quantification (LoQ);

- For yeast fungi Candida spp., opportunistic genital mycoplasmas Mycoplasma hominis, Ureaplasma urealyticum, Ureaplasma parvum the result can be either:

  - negative (“Not detected”; “-”);

  - positive, as “<4.0” Lg GE/mL, if the DNA concentration is below the limit of quantification (LoQ);

  - positive, as “X” Lg GE/mL, if the DNA concentration is above the limit of quantification (LoQ).

- To assess the composition of the microbiota, the proportion of the bacterial analyte (BA) relative to the total number of bacteria (TBL) is calculated using the formula BAP = 〖10〗Lg(BA) /〖10〗Lg(TBL) х 100%, where Lg (BA) and Lg (TBL) are the Lg values obtained for the target bacterial analyte and TBL, respectively.

Using this formula, the proportion of bacterial groups relative to TBL can also be calculated. This includes the proportions of normobiota, aerobes (facultative anaerobic microorganisms) associated with vaginitis/cervicitis and anaerobes (obligate anaerobic microorganisms) associated with bacterial vaginosis. As a result, the dominant microorganism/group of microorganisms can be identified.

- For pathogens, the result is considered valid regardless of the TBL and HGD values. It can be either:

  - negative (“Not detected”; “-”);

  - positive, as “<3.0” Lg GE/mL, if the DNA concentration is below the limit of quantification (LoQ);

  - positive, as “X” Lg GE/mL, if the DNA concentration is above the limit of quantification (LoQ).

| V. RESULTS AND PERFORMANCE |  |  |
| --- | --- | --- |
| 22. | Article 20.4.1(q) | Analytical Performance Data |
| 23. | Article 20.4.1(q) | Clinical Performance Data |

a.	Analytical specificity

For biomaterial samples containing DNA of the target analytes, the software of the detecting thermal cycler records positive results for specific amplification products in the corresponding tubes on the specified detection channels during amplification.

For biomaterial samples that do not contain DNA of the target analytes, the software of the detecting thermal cycler records negative results for specific amplification products in the corresponding tubes on the specified detection channels and positive amplification results for the internal control on the Hex detection channel (except for tubes without IC, see Table 1).

Analytical specificity of Femobiome II Multiplex qPCR Kit was evaluated against a panel of microorganisms detectable by this reagent kit and those potentially present in biomaterial sample. The panel included microorganisms in concentrations of at least 1.0х106 copies/mL. Absence of nonspecific positive amplification results and nonspecific cross-reactions of each component in the reagent kit against another target has been confirmed.

Table 9. Microorganisms used to assess analytical specificity

| Microorganisms |  |  |
| --- | --- | --- |
| Anaerococcus octavius | Haemophilus parainfluenzae | Mycoplasma hominis |
| Anaerococcus vaginalis | HSV-1 | Mycoplasma genitalium |
| Bacteroides fragilis | HSV-2 | Neisseria gonorrhoeae |
| Bacteroides ovatus | HPV 16 | Parvimonas micra |
| Bifidobacterium bifidum | HPV 18 | Peptoniphilus asaccharolyticus |
| Bifidobacterium breve | HPV 31 | Peptoniphilus septimus |
| Burkholderia cepacia | HPV 33 | Peptostreptococcus anaerobius |
| Burkholderia cenocepacia | HPV 35 | Peptostreptococcus faecalis |
| BVAB1 | HPV 39 | Porphyromonas vaginalis |
| BVAB2 | HPV 45 | Porphyromonas gingivalis |
| BVAB3 | HPV 51 | Prevotella copri |
| Candida albicans | HPV 52 | Prevotella stercorea |
| Candida parapsilosis | HPV 56 | Propionibacterium acnes |
| Chlamydia trachomatis | HPV 58 | Propionibacterium avidum |
| Corynebacterium glucuronolyticum | HPV 59 | Proteus mirabilis |
| Corynebacterium vaginale | HPV 66 | Pseudomonas aeruginosa |
| CMV | HPV 68 | Pseudomonas alcaligenes |
| Dialister invisus | Human herpesvirus 6 | Sneathia vaginalis |
| Eggerthella lenta | Human herpesvirus 8 | Staphylococcus saprophyticus |
| Eggerthella hominis | Lactobacillus salivarius | Staphylococcus aureus |
| Escherichia coli | Lactobacillus crispatus | Streptococcus pneumonia |
| Enterobacter intestinihominis | Lactobacillus iners | Streptococcus agalactiae |
| Enterococcus faecalis | Lactobacillus jensenii | Treponema pallidum |
| Enterococcus faecium | Lactobacillus mulieris | Trichomonas vaginalis |
| Epstein-Barr virus | Lactobacillus gasseri | Ureaplasma urealyticum |
| Fannyhessea vaginae | Leptotrichia massiliensis | Ureaplasma parvum |
| Fusobacterium nucleatum | Megasphaera cerevisiae | Veillonella dispar |
| Gardnerella vaginalis | Mobiluncus curtisii |  |
| Haemophilus influenzae | Mobiluncus mulieris |  |

b.	Analytical sensitivity (limit of detection)

The limit of detection (LoD) of the DNA targets detectable by this reagent kit was established by analyzing limiting dilutions of quantitatively characterized reference controls containing vectors with specific DNA sequences of the target analytes (controlled for the copy number of target genes). The LoD was also confirmed by the results obtained on dilutions of quantitatively characterized microorganism cultures. Analytical sensitivity for different DNA targets is provided in Table 10.

WARNING! The LoD depends on the elution volume of the DNA extraction kits (the larger the elution volume, the higher the LoD, which is due to a decrease in DNA concentration).

Table 10. Analytical sensitivity for different DNA targets.

| DNA targets | Analytical sensitivity in GE/mL |
| --- | --- |
| Trichomonas vaginalis | 1.0х102 |
| Candida spp., Candida albicans, Chlamydia trachomatis | 5.0х102 |
| TBL, HGD, Gardnerella vaginalis, Lactobacillus spp., Lactobacillus crispatus, Lactobacillus iners, Lactobacillus jensenii/mulieris, Lactobacillus gasseri/Lactobacillus paragasseri, Staphylococcus spp., Streptococcus spp., Streptococcus agalactiae, Mobiluncus spp., Haemophilus spp., Anaerococcus spp., Fannyhessea vaginae, Bacteroides spp., Porphyromonas spp., Prevotella spp., Enterobacteiaceae, Enterococcus spp., Sneathia spp., Leptotrichia spp., Fusobacterium spp., Megasphaera spp., Veillonella spp., Dialister spp., BVAB1, BVAB2, BVAB3, Peptostreptococcus spp., Bifidobacterium spp., Mycoplasma hominis, Ureaplasma urealyticum, Ureaplasma parvum, Neisseria gonorrhoeae, Mycoplasma genitalium, HSV-1, HSV-2, CMV | 1.0х103 |
| HPV 16, HPV 18, HPV 45, HPV 31/33/35/39/51/52/ 56/58/59/66/68 | 5.0х103 |

c.	Quantification range, limi

Reagent kit analytes quantification range is from 1.0x104 to 1.0x109 GE/mL. It was established by analyzing limiting dilutions of quantitatively characterized reference controls containing vectors with specific DNA sequences of the detectable analytes (controlled for the copy number of target genes). The quantification range was confirmed by the results obtained on dilutions of quantitatively characterized microorganism cultures.

The limit of quantification of this reagent kit is the lower limit of the linear range of the amplification, which corresponds to a value of 1.0x104 for all target analytes for the kit.

The coefficient of variation for the limit of quantification is no more than 10%.

d.	Diagnostic characteristics

The values of diagnostic sensitivity and specificity of the kit at 95% confidence level are provided in Table 11.

Table 11. Diagnostic sensitivity and specificity values.

| Biological material | Target analyte | Number of samples |  | Diagnostic sensitivity, % | Diagnostic specificity, % |
| --- | --- | --- | --- | --- | --- |
|  |  | Positive | Negative |  |  |
|  | Chlamydia | 100 | 300 | 100% (97.2-100) | 100% (99.1-100) |
|  | trachomatis |  |  |  |  |
|  | Neisseria gonorrhoeae | 100 | 300 | 100% (97.2-100) | 100% (99.1-100) |
| Epithelial swab from | Mycoplasma genitalium | 100 | 300 | 100% (97.2-100) | 100% (99.1-100) |
|  | Trichomonas vaginalis | 100 | 300 | 100% (97.2-100) | 100% (99.1-100) |
| the mucous |  |  |  |  |  |
|  | HSV-1 | 100 | 200 | 100% (97.2-100) | 100% (98.6-100) |
| membrane |  |  |  |  |  |
|  | HSV-2 | 100 | 200 | 100% (97.2-100) | 100% (98.6-100) |
| of the |  |  |  |  |  |
|  | CMV | 100 | 200 | 100% (97.2-100) | 100% (98.6-100) |
| cervical |  |  |  |  |  |
|  | HPV 16 | 100 | 200 | 100% (97.2-100) | 100% (98.6-100) |
| canal, vagina |  |  |  |  |  |
|  | HPV 18 | 100 | 200 | 100% (97.2-100) | 100% (98.6-100) |
|  | HPV 45 | 100 | 200 | 100% (97.2-100) | 100% (98.6-100) |
|  | HPV 31/33/35/39/51/ | 100 | 200 | 100% (97.2-100) | 100% (98.6-100) |
|  | 52/56/58/59/66/68 |  |  |  |  |

Diagnostic characteristics of the kit for quantitative DNA-based detection of normobiota representatives (Lactobacillus spp., Lactobacillus iners, Lactobacillus crispatus, Lactobacillus jensenii/mulieris, Lactobacillus gasseri/Lactobacillus paragasseri, Bifidobacterium spp.), quantitative DNA-based detection of opportunistic microorganisms (Ureaplasma urealyticum, Ureaplasma parvum, Mycoplasma hominis, Candida spp., Candida albicans, Staphylococcus spp., Streptococcus spp., Streptococcus agalactiae, Enterobacteriaceae, Enterococcus spp., Haemophilus spp., Gardnerella vaginalis, Fannyhessea vaginae, Mobiluncus spp., Anaerococcus spp., Peptostreptococcus spp., Bacteroides spp./Porphyromonas spp./Prevotella spp., Sneathia spp./Leptotrichia spp./Fusobacterium spp., Megasphaera spp./Veilonella spp./Dialister spp., BVAB1/BVAB2/BVAB3) and human genomic DNA are not estimated since the abovementioned analytes may be present in reproductive tract of healthy women, so the software only displays them in quantitative format. The diagnostic characteristics are, therefore, not applicable for quantitative methods.

e.	Measurement precision (repeatability and reproducibility)

The repeatability and reproducibility of measurements of this kit were evaluated by testing model samples. Model samples are prepared using vaginal and cervical biological material as a matrix and adding known quantities of microorganisms to it. Four concentrations were used: the limit of detection, limit of quantification and two values within the linear range of the amplification curve – 1.0x106 and 1.0x109. Each sample underwent all stages of the assay (DNA extraction, amplification and results detection).

The repeatability was evaluated at each of the four concentrations, in each of the packages (S, A, A-TL), using one instrument (DTprime 5M* for package S, DTprime 5X* for packages A, A-TL), by one operator, over a period of 5 days, with 2 runs per day and two repeats per run. The same reagent kit batch for each package was used for all runs.

The coefficient of variation for the repeatability of measurements for the “limit of detection” concentration was no more than 20%, for all three concentrations in the linear range – no more than 10% for all packages.

The reproducibility of measurements was evaluated at each of the four concentrations, in each of the packages (S, A, A-TL). For package S, two types of instruments (DTprime 5M*, DTlite 5S*) - two of each type, for packages A, A-TL - three instruments of one type (DTprime 5X*) were used. Two batches of reagent kits of each package were used for all runs. All runs were performed by 2 operators on 2 different days.

The coefficient of variation for reproducibility of measurements for the “limit of detection” concentration was no more than 20%, for all three concentrations in the linear range – no more than 10% for all packages.

| V. RESULTS AND PERFORMANCE |  |  |
| --- | --- | --- |
| 24. | Article 20.4.1(m) | Retesting Requirements |

|  | Result | Possible cause | Solution |
| --- | --- | --- | --- |
| C+ | - | Operation error PCR inhibition Violation of storage and handling requirements | Repeat the entire test Dispose of the current batch |
| C- | + | Contamination | Dispose of the current batch Perform decontamination procedures |
| IC | Invalid | PCR inhibition | Repeat the entire test Re-take sample |

| VI. ADDITIONAL INFORMATION |  |  |
| --- | --- | --- |
| 25. | Article 20.4.1(t) | Disposal of Used Materials |

| VI. ADDITIONAL INFORMATION |  |  |
| --- | --- | --- |
| 26. | Article 20.4.1(x) | Revision Date of Instructions for Use |

| VI. ADDITIONAL INFORMATION |  |  |
| --- | --- | --- |
| 27. | Article 20.4.1(v) | Standards with which Compliance is Claimed |

| VI. ADDITIONAL INFORMATION |  |  |
| --- | --- | --- |
| 28. | Article 20.4.1(r) | Check Before Use |

| VI. ADDITIONAL INFORMATION |  |
| --- | --- |
| 29. | Technical Support Contact |

VII. SYMBOLS/PICTOGRAMS

|  | Temperature limit |  | Date of manufacture |
| --- | --- | --- | --- |
|  | Contains sufficient for <n> tests |  | Consult instructions for use |
|  | Use-by date |  | Catalogue number |
|  | Batch code |  | Manufacturer |
|  | Caution |  | Keep away from sunlight |
|  | Do not reuse |  | Non-sterile |

Annex A

Shortened procedure for DNA extraction from epithelial swabs from the mucous membrane of cervical canal and vagina using PREP-NA PLUS reagent kit

WARNING!

	Before starting work:

•	heat the thermostat to 65 °С;

•	take the PREP-NA PLUS reagent kit out of the refrigerator and check that there is no precipitate in the lysis solution. If precipitate is observed, heat the vial with the lysis solution on a thermostat preheated to 65°C to dissolve the precipitate completely. Then stir the solution by turning the vial upside down 5-10 times (avoiding foaming). Before use, cool the solution to room temperature (18°C to 25°C). The precipitate can also be dissolved at room temperature (18°C to 25°C) for approximately 12 hours.

	Caps may open during heating! Use tubes with self-locking caps or programmable thermostats with a clamp lid.

Procedure:

1.	Mark a 1.5 mL plastic tube as negative control (C-).

2.	Enter 300 μL of lysis solution into each marked tube with 100 μL of biomaterial and into the “C-

tube. Do not touch the edges of the tube.

3.	Add 100 μL of negative control into the “C-“ tube.

4.	Close the tubes tightly and shake on vortex for 3-5 seconds.

5.	Heat the tubes on thermostat at 65 °С for 5 minutes.

6.	Spin the tubes on vortex for 3-5 seconds.

7.	Add 400 μL of precipitation buffer solution into each tube without touching the edges of the tube, close the tubes and shake on vortex for 3-5 seconds.

8.	Centrifuge the tubes at RCF(g)  12,000-16,000  at  room  temperature  (from  18°С  to  25°С)  for 10 minutes.

9.	Using a separate tip for each tube, carefully remove as much supernatant as possible without disturbing the precipitate.

10.	Add 500 μL of wash solution No. 1 to the precipitate, close the tubes and mix the contents by carefully turning the tubes upside down 3-5 times.

11.	Centrifuge the  tubes  at  RCF(g)  12,000-16,000  at  room  temperature  (from  18°С  to  25°С)  for 1 minute.

12.	Using a separate tip for each tube, carefully remove as much supernatant as possible without disturbing the precipitate.

13.	Add 300 μL of wash solution No. 2 to the precipitate, close the tubes and mix the contents by carefully turning the tubes upside down 3-5 times.

14.	Centrifuge the  tubes  at  RCF(g)  12,000-16,000  at  room  temperature  (from  18°С  to  25°С)  for 1 minute.

15.	Using a separate tip for each tube, carefully remove as much supernatant as possible without disturbing the precipitate. It is acceptable to leave up to 20-30 μL of liquid above the precipitate.

16.	Open the tubes and dry the precipitate at 65 °С for 5 minutes.

17.	Add 300 μL of dilution buffer to the precipitate, shake the tubes on vortex for 3-5 seconds and centrifuge at RCF(g) 1,000-3,000 for 3-5 seconds.

18.	Heat the tubes on thermostat at 65 °С for 5 minutes. Shake the tubes on vortex for 3-5 seconds.

19.	Centrifuge the tubes at RCF(g) 12,000-16,000 for 30 seconds at room temperature (from 18°С to 25°С) to spin down the condensate.

The DNA sample can now be introduced into PCR mix.

Before using the DNA sample for PCR, thaw the DNA sample and the negative control at room temperature (from 18°С to 25°С) or in a refrigerator (from 2°С to 8°С), then shake the tubes with the DNA sample and the negative control on vortex for 3-5 seconds, then centrifuge at RCF(g) 1,000-3,000 for   1-3 seconds.

WARNING! The DNA sample can be thawed only once!

# REVISION HISTORY

| Revision No. | Rev. Date | Description |
| --- | --- | --- |
| 00 | 14.11.2025 | First Issue |

# DOCUMENT APPROVAL

| Prepared By | Approved By |
| --- | --- |
| Quality Management Representative | General Manager |
| Elife ÖZDEMİR | Oktay ÖZTÜRK |