PRODUCT NAME:

Femobiome PrimaScreen Multiplex qPCR Kit

INTENDED USE:

Femobiome PrimaScreen Multiplex qPCR Kit designed for the quantitative determination of DNA from normobiota representatives (Lactobacillus iners and Lactobacillus non-iners) and opportunistic in biological specimens obtained from the female reproductive tract (vaginal and cervical mucosal epithelial swabs/scrapings) via real-time polymerase chain reaction.

Headquarters: Neighborhood of Dumlupınar, Nur St. No: 7/3 Post Code: 41250 Kartepe/Kocaeli/TURKEY

Branch: TUBITAK MAM Technology Free Zone, Neighborhood of Barıs SB, 5002 Street, New Tech. Building, Blok A, No: 4A/ Z09 Post Code: 41400

Gebze/Kocaeli/TURKEY

Trd Register Number: 34218

Email: info@genomer.com.tr

https://genomer.com.tr

Purpose

This document constitutes the Instructions for Use (IFU) for the Femobiome PrimaScreen Multiplex qPCR Kit, an In Vitro Diagnostic (IVD) medical device.

The European Union's In Vitro Diagnostic Medical Devices Regulation (IVDR) 2017/746 aims to ensure a high level of health and safety protection for patients and users throughout the entire lifecycle of the device.

These Instructions for Use have been prepared to fully satisfy all obligations stipulated under Annex I: General Safety and Performance Requirements of the IVDR, specifically within Chapter III: Information Supplied by the Manufacturer (Article 20).

The primary objective is to ensure the safe, accurate, and effective use of the device for its intended purpose. These instructions encompass all necessary technical, procedural, and legal information for laboratory professionals, including device performance, limitations, potential risks, and appropriate sample management.

Regulatory Compliance and Transparency

This guide not only employs pictograms and symbols strictly in accordance with national and international standards (specifically IVDR, EN ISO 15223-1, and EN ISO 18113) but also transparently presents all legally required device identifiers (UDI, Manufacturer, Authorized Representative).

Warning: It is imperative to read, understand, and meticulously adhere to these instructions in their entirety prior to using the device to ensure safe and accurate results. Failure to comply may lead to erroneous results, potential hazards, or invalidation of the CE marking.

| I. DESCRIPTIVE INFORMATION |  |  |  |
| --- | --- | --- | --- |
| 1. | Device Name, Trade Name, and Model | Article 20.3(a), 20.4.1(b) | Femobiome PrimaScreen Multiplex qPCR Kit |
| 2. | In Vitro Diagnostic Device | Article 20.3(b) | Femobiome PrimaScreen Multiplex qPCR Kit designed for the quantitative determination of DNA from normobiota representatives (Lactobacillus iners and Lactobacillus non-iners) and opportunistic microorganisms in biological specimens obtained from the female reproductive tract (vaginal and cervical mucosal epithelial swabs/scrapings) via real-time polymerase chain reaction. |
| 3. | Manufacturer Information | Article 20.3(d) | GENOMER BIOTECHNOLOGY R&D IND. TRD. LIMITED COMPANY, Headquarters: Neighborhood of Dumlupınar, Nur St. No: 7/3 Post Code: 41250 Kartepe/Kocaeli/TURKEY Branch: TUBITAK MAM Technology Free Zone, Neighborhood of Barıs SB, 5002 Street, New Tech. Building, Blok A, No: 4A/ Z09 Post Code: 41400 Gebze/Kocaeli/TURKEY Trd Register Number: 34218 Email: info@genomer.com.tr Web:https://genomer.com.tr |
| 4. | Authorized Representative Information | Article 20.3(e) | Required only if the manufacturer is located outside the EU. |
| 5. | Lot/Batch Number | Article 20.3(f) | UL- FBPS-XX |
| 6. | Expiry/Use-by Date | Article 20.3(g) | YYYY/AA/GG |
| 7. | Catalog Number | Article 20.3(c) | M-FBPS-0601v01-24 |
| 8. | Unique Device Identifier (UDI) Reference | Article 20.3(n) | 868512670MFBASKB |

The results obtained in accordance with the procedures and analytical acceptance criteria specified in the aforementioned Instructions for Use (IFU) guide confirm the conformity of the Femobiome PrimaScreen Multiplex qPCR Kit device with the General Safety and Performance Requirements (Annex I).

| II. INTENDED USE, SCOPE, AND PRINCIPLE |  |  |  |
| --- | --- | --- | --- |
| 9. | Intended Use and Function | Article 20.4.1(c) | INTENDED USE |

Reagent kit for the quantitative assessment of normobiota, bacterial vaginosis pathogens, opportunistic mycoplasmas, and Candida species in the female reproductive tract using real-time PCR (Femobiome PrimaScreen Multiplex qPCR Kit); hereinafter referred to as the "reagent kit."

Intended Use: The reagent kit is designed for the quantitative determination of DNA from normobiota representatives (Lactobacillus iners and Lactobacillus non-iners) and opportunistic microorganisms (Ureaplasma urealyticum, Ureaplasma parvum, Mycoplasma hominis, Candida spp., Candida albicans, Gardnerella vaginalis, Fannyhessea vaginae, Mobiluncus spp., Anaerococcus spp./ Peptostreptococcus spp., Bacteroides spp./ Porphyromonas spp./ Prevotella spp., Sneathia spp./ Leptotrichia spp./ Fusobacterium spp., Megasphaera spp./ Veillonella spp./ Dialister spp., BVAB1/ BVAB2/ BVAB3) in biological specimens obtained from the female reproductive tract (vaginal and cervical mucosal epithelial swabs/scrapings) via real-time polymerase chain reaction.

Functional Purpose: In vitro diagnostics (IVD).

Indications for Analysis:

- Infectious and inflammatory diseases of the reproductive tract;

- Reproductive health disorders (infertility, recurrent pregnancy loss). There are no known contraindications for use.

Population and Demographic Aspects: The reagent kit is intended for the examination of female individuals who have completed the pubertal period (post-pubertal).

Scope of Application: The reagent kit is suitable for use in clinical diagnostic laboratories of medical institutions.

Potential Users: Qualified personnel trained in molecular diagnostic methods and clinical laboratory safety protocols, including clinical diagnostic laboratory physicians and laboratory technicians (medical laboratory technologists).

Usage Requirement: The reagent kit must be used strictly for its intended purpose in accordance with these instructions for use.

| II. INTENDED USE, SCOPE, AND PRINCIPLE |  |  |  |
| --- | --- | --- | --- |
| 10. | Test Principle | Article 20.4.1(e) | METHOD |

The principle of the method is based on the process of DNA amplification, which consists of repeating cycles of thermal DNA denaturation, annealing of primers to complementary sequences, and subsequent extension of polynucleotide chains by Taq polymerase.

To enhance the sensitivity and specificity of the reaction, a "hot start" is provided. The "hot start" is ensured by the preparation method of the reaction mixture, which consists of two layers separated by a paraffin interlayer. The mixing of the layers and their transformation into a reaction mixture occurs only after the paraffin melts, which eliminates non-specific binding of primers to the target DNA at lower temperatures.

DNA probes, each containing a fluorescent label and a fluorescence quencher, are added to the amplification mixture. Upon the formation of a specific product, the DNA probe is degraded, and the effect of the quencher on the fluorescent label ceases, leading to an increase in the fluorescence level, which is recorded by the detecting thermal cycler. The number of degraded probes (and, consequently, the fluorescence level) increases proportionally to the amount of specific amplification products formed. The fluorescence level is measured at each amplification cycle in real-time.

The reagent kit includes amplification mixtures specific for the detection of Lactobacillus DNA and opportunistic microorganisms, as well as for the determination of auxiliary analytes: DNA of all bacteria (TBC – Total Bacterial Count) and human genomic DNA (HGD).

An Internal Control (IC) is included in the amplification mixtures (with the exception of tubes No. 1 and No. 7) to assess the quality of the polymerase chain reaction in each individual tube.

To verify the orientation of the strip in the thermal block of the real-time PCR instrument, a fluorescently labeled oligonucleotide (Cy5.5 dye), designated as the "Marker," is added to the amplification mixture in tube No. 2. It is used by the instrument to determine the position of the strip in the thermal block of the detecting thermal cycler. After the amplification run, the software compares the operator-defined marker position with its actual location. If a discrepancy is detected (due to incorrect strip orientation), it alerts the operator to this mismatch. In such cases, it is necessary to check the placement of the strip in the thermal block (the first tube of the strip is marked with a blue buffer) and correct the tube identifiers in the protocol.

The DNA probes used for the detection of the amplification products of the target microorganism genome fragments include the following fluorescent labels: Fam, Hex, Rox, Cy5, and Cy5.5. The DNA probe used for the detection of the internal control (IC) amplification product contains the Hex fluorescent dye.

The use of multiple fluorescent dyes allows for a reduction in the number of tubes and the volume of biological material required for the study, as it enables the simultaneous registration of results from different amplification reactions occurring within a single tube.

Table 1 lists the detectable parameters, color coding, and detection channels for the amplification products.

| Test tube number | Dye label/detection channel |  |  |  |  | Color coding of the amplification mixture |
| --- | --- | --- | --- | --- | --- | --- |
|  | Fam | Hex | Rox | Cy5 | Cy5.5 |  |
| 1 | Lactobacillus не-iners | Lactobacillus iners | – | ОКБ | Gardnerella vaginalis | Blue |
| 2 | Bacteroides spp./ Porphyromonas spp./ Prevotella spp. | ВК | – | Mobiluncus spp. | Маркер | Colorless or pink |
| 3 | Anaerococcus spp. | ВК | Fannyhessea vaginae | – | – |  |
| 4 | Sneathia spp./ Leptotrichia spp./ Fusobacterium spp. | ВК | – | – | – |  |
| 5 | Megasphaera spp./ Veillonella spp./ Dialister spp. | ВК | – | – | – |  |
| 6 | BVAB1/ BVAB2/ BVAB3 | ВК | – | – | – |  |
| 7 | Mycoplasma hominis | Ureaplasma urealyticum | Ureaplasma parvum | ГДЧ | – |  |
| 8 | Candida spp. | ВК | Candida albicans | Peptostreptococcus spp. | – |  |

Table 1: Detectable indicators, color coding and detection channels of amplification products

Determination of the decimal logarithm (Lg) of the concentration (number of target DNA copies in 1.0 ml of DNA preparation) is performed using the threshold cycle comparison method, also known as the ΔΔCt method. The advantage of the ΔΔCt method is the determination of the amount of target DNA even under conditions of suboptimal PCR efficiency without the need to construct a calibration curve (Shefe J. H. Quantitative real-time RT-PCR data analysis: current concepts and the novel «gene expression’s Ct difference» formula [Text] / J. H. Shefe, K. E. Lehmann, I. R. Buschmann, T. Unger, H. Funke-Kaiser // J. Mol. Med. – 2006. – Vol. 84. – P. 901–910 [DOI 10.1007/s00109-006-0097-6]).1

The calculation takes into account the values of the threshold cycle2 (Ct) and amplification efficiency. The threshold cycle is the point on the abscissa axis of the product accumulation graph relative to the amplification cycle number. To find the threshold cycle, the Crossing point method is used—determination of the second derivative maximum (Cp) (Rebrikov et al., 2014). PCR efficiency is a numerical value indicating how many times the number of DNA fragments changes during one amplification cycle.

The calculation is performed according to the formula:

𝐿𝑜𝑔(𝑁𝑂a) = 𝐿𝑜𝑔(𝐸) × (𝐶𝑝 − 𝐶𝑝a),	(1)

where:

N0a – initial number of target DNA copies in 1.0 ml of the DNA preparation of the analyzed sample;

E – PCR efficiency;

Cp – threshold cycle for a sample containing a single DNA molecule;

Cpa – threshold cycle for the analyzed sample.

The study consists of the following stages: DNA extraction (sample preparation) and real-time PCR amplification of DNA with result detection using the FEMOFLOR® AlphaScreen reagent kit.

| II. INTENDED USE, SCOPE, AND PRINCIPLE |  |  |  |
| --- | --- | --- | --- |
| 11. | Adequate Number of Tests | Article 20.3(k) | CONTENTS |

The Femobiome PrimaScreen Multiplex qPCR Kit contains PCR-mix, Taq-polymerase solution, mineral oil and positive control sample. The detailed description of content is represented in Table.

| Reagent | Description | Amount | Total volume |
| --- | --- | --- | --- |
| Paraffin sealed PCR-mix | Colorless pink, or blue liquid under white waxy fraction | 24 strips of 8 tubes | 20 µL in each tube |
| Taq-polymerase solution | Colorless transparent liquid | 2 tubes | по 1,0 мл |
| Mineral oil | Colorless transparent viscous oily liquid | 4 tubes | 1.0 mL |
| Positive control* | Colorless transparent liquid | 1 tube | 160 мкл |
| Strip’s caps | 24 caps |  |  |

Table 2: Femobiome PrimaScreen Multiplex qPCR Kit content

All components of the reagent kit are ready-to-use and do not require additional preparation.

The reagent kit is designed for 24 determinations (no more than 8 runs), including the analysis of unknown samples, negative control samples, and positive control samples.

| III. SAFETY AND PRECAUTIONS |  |  |  |
| --- | --- | --- | --- |
| 12. | Warnings and Precautions | Article 20.4.1(n) | All necessary general warnings for safe use |
| 13. | Biohazard Warning | Article 20.4.1(p) | Specimens pose a potential infection risk |
| 14. | Hazardous Chemical Substances | Article 20.4.1(o) | If kit components contain hazardous substances |
| 15. | Restriction on Use | Article 20.1(f) | For use only by trained laboratory personnel |

| III. GÜVENLİK VE ÖNLEMLER |  |  |  |
| --- | --- | --- | --- |
| 16. | Special Storage Conditions | Article 20.4.1(j) | Required temperature range for reagent storage |

| IV. TEST PROCEDURE AND QUALITY CONTROL |  |  |  |
| --- | --- | --- | --- |
| 17. | Sample Type | Article 20.4.1(d) | Type of specimen accepted for testing |

| IV. TEST PROCEDURE AND QUALITY CONTROL |  |  |  |
| --- | --- | --- | --- |
| 18. | Refer to Instructions for Use | Article 20.3(j) | Must refer to instructions for detailed procedure |
| 19. | Procedure Steps | Article 20.4.1(l) | Flowchart from nucleic acid extraction to qPCR cycling |

| IV. TEST PROCEDURE AND QUALITY CONTROL |  |  |  |
| --- | --- | --- | --- |
| 20. | Quality Control (QC) Procedure | Article 20.4.1(m) | Details for use of Positive, Negative and Internal Control (IC) |

| V. RESULTS AND PERFORMANCE |  |  |  |
| --- | --- | --- | --- |
| 21. | Interpretation of Results | Article 20.4.1(m) | Ct cut-off values and result table for each pathogen |

| V. RESULTS AND PERFORMANCE |  |  |  |
| --- | --- | --- | --- |
| 22. | Analytical Performance Data | Article 20.4.1(q) | Technical parameters such as LOD, Specificity, Precision, Linear Range. |
| 23. | Clinical Performance Data | Article 20.4.1(q) | Diagnostic Sensitivity, Specificity, and expected values (prevalence) in pediatric population. |

a. Analytical Specificity In human biological material samples containing the DNA of the analytes being detected, the software of the detecting thermal cycler must record positive amplification results for specific products in the corresponding tubes via the designated detection channels during the amplification process. In human biological material samples that do not contain the DNA of the analytes being detected, the software of the detecting thermal cycler must record negative amplification results for specific products in the corresponding tubes via the designated detection channels and a positive amplification result for the internal control via the Hex detection channel (with the exception of tubes without IC, see Table 1).The analytical specificity of the FEMOFLOR® AlphaScreen kit was evaluated using a panel of microorganisms detectable by the reagent kit and microorganisms potentially present in the biomaterial sample, with a microorganism concentration of at least $1.0 \times 10^6$ copies/ml. Testing results demonstrated the absence of non-specific positive amplification results.The absence of cross-reactive non-specific reactions for each component included in the reagent kit with respect to other targets within the system was demonstrated.

| Names of microorganisms |  |  |
| --- | --- | --- |
| Anaerococcus vaginalis | Fusobacterium nucleatum | Mycoplasma genitalium |
| Bacteroides fragilis | Gardnerella vaginalis | Neisseria gonorrhoeae |
| Bifidobacterium bifidum | Haemophilus influenzae | Parvimonas micra |
| Burkholderia cepacia | HSV1 | Peptoniphilus asaccharolyticus |
| BVAB1 | HSV2 | Peptostreptococcus anaerobius |
| BVAB2 | HPV 16 | Porphyromonas vaginalis |
| BVAB3 | HPV 18 | Prevotella copri |
| Candida albicans | HPV 45 | Propionibacterium acnes |
| Candida parapsilosis | Human herpesvirus 6 | Proteus mirabilis |
| Chlamydia trachomatis | Human herpesvirus 8 | Pseudomonas aeruginosa |
| Corynebacterium vaginale | Lactobacillus crispatus | Sneathia vaginalis |
| CMV | Lactobacillus iners | Staphylococcus aureus |
| Dialister invisus | Lactobacillus jensenii | Streptococcus agalactiae |
| Eggerthella lenta | Lactobacillus mulieris | Treponema pallidum |
| Escherichia coli | Lactobacillus gasseri | Trichomonas vaginalis |
| Enterobacter intestinihominis | Leptotrichia massiliensis | Ureaplasma urealyticum |
| Enterococcus faecalis | Megasphaera cerevisiae | Ureaplasma parvum |
| Epstein-Barr virus | Mobiluncus curtisii |  |
| Fannyhessea vaginae | Mycoplasma hominis |  |

| V. SONUÇLAR VE PERFORMANS |  |  |  |
| --- | --- | --- | --- |
| 24. | Retesting Requirements | Article 20.4.1(m) | How to repeat invalid or equivocal results. |

| VI. ADDITIONAL INFORMATION |  |  |  |
| --- | --- | --- | --- |
| 25. | Disposal of Used Materials | Article 20.4.1(t) | Instructions for medical and hazardous waste management |
| 26. | Revision Date of Instructions for Use | Article 20.4.1(x) | 2025-11-25 |
| 27. | Standards with which Compliance is Claimed | Article 20.4.1(v) | ISO 15223-1, ISO 18113 Reference to relevant standards such |
| 28. | Check Before Use | Article 20.4.1(r) | Actions in case of damaged packaging or reagent deterioration |
| 29. | Technical Support Contact | - | Contact details for questions and technical issues. |

VII. SYMBOLS/PICTOGRAMS

# REVISION HISTORY

| Revision No. | Rev. Date | Description |
| --- | --- | --- |
| 00 | 14.11.2025 | First Issue |

# DOCUMENT APPROVAL

| Prepared By | Approved By |
| --- | --- |
| Quality Management Representative | General Manager |
| Elife ÖZDEMİR | Oktay ÖZTÜRK |