PRODUCT NAME:

PREP-MB MAX DNA Extraction Kit

INTENDED USE:

The PREP-MB MAX DNA Extraction Kit is intended for human, bacterial, viral, and fungal DNA extraction from human biological material (whole peripheral blood; smears/scrapings from urogenital tract and rectum; urine; ejaculate; milk; faeces) for further PCR analysis.

This medical device is an auxiliary agent in clinical laboratory diagnostics.

Headquarters: Neighborhood of Dumlupınar, Nur St. No: 7/3 Post Code: 41250 Kartepe/Kocaeli/TURKEY

Branch: TUBITAK MAM Technology Free Zone, Neighborhood of Barıs SB, 5002 Street, New Tech. Building, Blok A, No: 4A/ Z09 Post Code: 41400

Gebze/Kocaeli/TURKEY

Trd Register Number: 34218

Email: info@genomer.com.tr

https://genomer.com.tr

Purpose

This document constitutes the Instructions for Use (IFU) for the PREP-MB MAX DNA Extraction Kit, an In Vitro Diagnostic (IVD) medical device.

The European Union's In Vitro Diagnostic Medical Devices Regulation (IVDR) 2017/746 aims to ensure a high level of health and safety protection for patients and users throughout the entire lifecycle of the device.

These Instructions for Use have been prepared to fully satisfy all obligations stipulated under Annex I: General Safety and Performance Requirements of the IVDR, specifically within Chapter III: Information Supplied by the Manufacturer (Article 20).

The primary objective is to ensure the safe, accurate, and effective use of the device for its intended purpose. These instructions encompass all necessary technical, procedural, and legal information for laboratory professionals, including device performance, limitations, potential risks, and appropriate sample management.

Regulatory Compliance and Transparency

This guide not only employs pictograms and symbols strictly in accordance with national and international standards (specifically IVDR, EN ISO 15223-1, and EN ISO 18113) but also transparently presents all legally required device identifiers (UDI, Manufacturer, Authorized Representative).

Warning: It is imperative to read, understand, and meticulously adhere to these instructions in their entirety prior to using the device to ensure safe and accurate results. Failure to comply may lead to erroneous results, potential hazards, or invalidation of the CE marking.

| I. DESCRIPTIVE INFORMATION |  |  |  |
| --- | --- | --- | --- |
| 1. | Article 20.3(a), 20.4.1(b) | Device Name, Trade Name, and Model | PREP-MB MAX DNA Extraction Kit |
| 2. | Article 20.3(b) | In Vitro Diagnostic Device | The PREP-MB MAX DNA Extraction Kit is intended for human, bacterial, viral, and fungal DNA extraction from human biological material (whole peripheral blood; smears/scrapings from urogenital tract and rectum; urine; ejaculate; milk; faeces) for further PCR analysis. This medical device is an auxiliary agent in clinical laboratory diagnostics. |
| 3. | Article 20.3(d) | Manufacturer Information | GENOMER BIOTECHNOLOGY R&D IND. TRD. LIMITED COMPANY, Headquarters: Neighborhood of Dumlupınar, Nur St. No: 7/3 Post Code: 41250 Kartepe/Kocaeli/TURKEY Branch: TUBITAK MAM Technology Free Zone, Neighborhood of Barıs SB, 5002 Street, New Tech. Building, Blok A, No: 4A/ Z09 Post Code: 41400 Gebze/Kocaeli/TURKEY Trd Register Number: 34218 Email: info@genomer.com.tr Web:https://genomer.com.tr |
| 4. | Article 20.3(e) | Authorized Representative Information | Required only if the manufacturer is located outside the EU. |
| 5. | Article 20.3(f) | Lot/Batch Number | UL- MBM-XX |
| 6. | Article 20.3(g) | Expiry/Use-by Date | YYYY/AA/GG |
| 7. | Article 20.3(c) | Catalog Number | I-MBM-v01-48 |
| 8. | Article 20.3(n) | Unique Device Identifier (UDI) Reference | 868512670IMBMMG |

The results obtained in accordance with the procedures and analytical acceptance criteria specified in the aforementioned Instructions for Use (IFU) guide confirm the conformity of the PREP-MB MAX DNA Extraction Kit device with the General Safety and Performance Requirements (Annex I).

| II. INTENDED USE, SCOPE, AND PRINCIPLE |  |  |
| --- | --- | --- |
| 9. | Article 20.4.1(c) | Intended Use and Function |

The PREP-MB MAX DNA Extraction Kit is intended for human, bacterial, viral, and fungal DNA extraction from human biological material (whole peripheral blood; smears/scrapings from urogenital tract and rectum; urine; ejaculate; milk; faeces) for further PCR analysis.

This medical device is an auxiliary agent in clinical laboratory diagnostics.

The application of the kit does not depend on population and demographic aspects. There are no contradictions for use of the PREP-MB MAX DNA Extraction Kit.

The PREP-MB MAX DNA Extraction Kit can be used in clinical and diagnostic laboratories of medical institutions and research practice.

Potential users: personnel qualified in molecular diagnostics methods and working in the clinical and diagnostic laboratory.

It is necessary to apply the kit only as directed in this instruction for use.

| II. INTENDED USE, SCOPE, AND PRINCIPLE |  |  |
| --- | --- | --- |
| 10. | Article 20.4.1(e) | Test Principle - Method |

Method: lysis and release of nucleic acids induced by a chaotropic agent (guanidine thiocyanate), followed by sorption on paramagnetic nanoparticles and washing out the impurities.

| II. INTENDED USE, SCOPE, AND PRINCIPLE |  |  |
| --- | --- | --- |
| 11. | Article 20.3(k) | Adequate Number of Tests |

The PREP-MB MAX DNA Extraction Kit content is represented in Tables 1.

| Reagent | Description | Total volume | Amount |
| --- | --- | --- | --- |
| Lysis solution | Colorless transparent foamy liquid | 7.2 mL | 1 vial |
| Sorbent | Brown suspension | 2.88 mL (1.44 mL in each) | 2 tubes |
| Binding solution | Colorless transparent liquid | 12 mL | 1 vial |
| Wash solution No. 1 | Colorless transparent liquid | 24 mL | 1 vial |
| Wash solution No. 2 | Colorless transparent liquid | 24 mL | 1 vial |
| Elution solution | Colorless transparent liquid | 14.4 mL | 1 vial |
| Negative control | Colorless transparent liquid | 2.4 mL (1.2 mL in each) | 2 tubes |
| Magnetic rods | Dark rods in polypropylene blister | 48 rods (16 rods in each blister) | 3 blisters |

All components are ready to use and do not require additional preparation for operation.

The reagent kit in package N is designed for DNA extraction from 48 test samples (including control samples).

REAGENTS AND EQUIPMENT REQUIRED BUT NOT PROVIDED

## Specimen collection

- Sterile single use swabs, single use sterile containers to collect clinical material;

- Sterile tubes containing transport media: “GENOMER” made STOR F Transport Medium, or equivalent for the transportation of the sample.

- For blood collection: 2.0 or 4.0 mL Vacuette blood collection tubes with anticoagulant, for example, salt of EDTA at a final concentration of 2.0 mg/mL or sodium citrate anticoagulant.

## NA extraction

- Biological safety cabinet class II-III;

- Refrigerator;

- High speed centrifuge (RCF(g) at least 16,000);

- Vortex mixer;

- Solid-state thermostat (temperature range 25-98 °C);

- Magnetic homogenizer;

- Single channel pipettes (dispensers covering 2.0-1000 μL volume range);

- RNase and DNase free filtered pipette tips (volume 20 μL, 200 μL, 1000 μL);

- 1.5 mL tubes (SSI-1260 are recommended);

- Tube rack for 1.5 mL tubes;

- Electric laboratory aspirator with trap flask for the removal of supernatant (for package N);

- RNase and DNase free non-filtered pipette tips for aspirator with trap flask (for package N);

- Specialized forceps or plastic forceps for arranging the magnetic rods (for package N);

- Physiological saline solution 0.9% NaCl (sterile);

- Container for used pipette tips, tubes and other consumables;

- Powder-free surgical gloves;

- Disinfectant solution.

| III. SAFETY AND PRECAUTIONS |  |  |
| --- | --- | --- |
| 12. | Article 20.4.1(n) | Warnings and Precautions |
| 13. | Article 20.4.1(p) | Biohazard Warning |
| 14. | Article 20.4.1(o) | Hazardous Chemical Substances |
| 15. | Article 20.1(f) | Restriction on Use |

Only personnel trained in the methods of molecular diagnostics and the rules of work in the clinical and diagnostic laboratory are allowed to work with the kit.

Handle and dispose all biological samples, reagents and materials used to carry out the assay as if they were able to transmit infective agents. The samples must be exclusively employed for certain type of analysis. Samples must be handled under a laminar flow hood. Tubes containing different samples must never be opened at the same time. Pipettes used to handle samples must be exclusively employed for this specific purpose. The pipettes must be of the positive dispensation type or be used with aerosol filter tips. The tips employed must be sterile, free from the DNases and RNases, free from DNA and RNA. The reagents must be handled under a laminar flow hood. The reagents required for amplification must be prepared in such a way that they can be used in a single session. Pipettes used to handle reagents must be exclusively employed for this specific purpose. The pipettes must be of the positive dispensation type or be used with aerosol filter tips. The tips employed must be sterile, free from the DNases and RNases, free from DNA and RNA. Avoid direct contact with the biological samples reagents and materials used to carry out the assay. Wear powder-free surgical gloves. Wear protective clothing (work clothes and personal protective equipment) working with microorganisms classified as particularly pathogenic. The protective clothing and personal protective equipment must comply with the work to be performed and health and safety requirements. Avoid producing spills or aerosol. Any material being exposed to biological samples must be treated for at least 30 minutes with disinfecting solution or autoclaved for 1 hour at 121°C before disposal.

Molecular biology procedures, such as nucleic acids extraction, reverse transcription, PCR-amplification and detection require qualified staff to avoid the risk of erroneous results, especially due to the degradation of nucleic acids contained in the samples or sample contamination by amplification products.

All the liquid solutions are designed for single use and can not be used more than once in amplification reactions. Plastic tubes do not contain phthalates. Do not breathe gas/fumes/vapor/spray produced by the components of the kit. Do not eat/drink components of the kit. Avoid contact with eyes. Only use the reagents provided in the kit and those recommended by manufacturer. Do not mix reagents from different batches. Do not use reagents from third party manufacturers’ kits. All laboratory equipment, including pipettes, test tube racks, laboratory glassware, lab coats, bouffant caps, etc., as well as reagents should be strictly stationary. It is not allowed to move them from one room to another. Equip separate areas for the extraction/preparation of amplification reactions and for the amplification/detection of amplification products. Never introduce an amplification product in the area designed for extraction/preparation of amplification reactions. Wear lab coats, gloves and tools, which are exclusively employed for the extraction/preparation of the amplification reaction and for the amplification/detection of the amplification products. Never transfer lab coats, gloves and tools from the area designed for amplification/detection of the amplification products to the area designed for extraction/preparation of amplification reactions. Remove waste materials (tubes, tips) only in a special closed container containing a disinfectant solution. Work surfaces, as well as rooms where NA extraction and PCR are performed, must be irradiated with bactericidal irradiators for 30 minutes before and after the work.

Waste materials are disposed of in accordance with local and national standards. All surfaces in the laboratory (work tables, test tube racks, equipment, etc.) must be treated daily with disinfecting solution.

Emergency actions

Eye Contact: If any component of this kit enters the eyes, wash eyes gently under potable running water for 15 minutes or longer, making sure that the eyelids are held open. If pain or irritation occurs, obtain medical attention.

Skin Contact: If any component of this kit contacts the skin and causes discomfort, remove any contaminated clothing. Wash affected area with plenty of soap and water. If pain or irritation occurs, obtain medical attention.

Ingestion: If any component of this kit is ingested, wash mouth out with water. If irritation or discomfort occurs, obtain medical attention.

| III. SAFETY AND PRECAUTIONS |  |  |
| --- | --- | --- |
| 16. | Article 20.4.1(j) | Special Storage Conditions |

Expiry date – 12 months from the date of production.

The PREP-MB MAX DNA Extraction Kit must be transported in thermoboxes with ice packs by all types of roofed transport at temperatures inside the thermoboxes corresponding to storage conditions of the kit components.

It is allowed to transport the kit in thermoboxes with ice packs by all types of roofed transport at temperatures from 2°C to 25°C inside the thermoboxes, but for no longer than 5 days.

Kits transported with violation of temperature conditions must not be used.

All components of the PREP-MB MAX DNA Extraction Kit must be stored in a refrigerator or cooling chamber at temperatures from 2°C to 8°C over the storage period.

A little precipitate is allowed in lysis solution, binding solution and wash solution No.1.

Cartridges, lysis solution, binding solution and wash solution No.1 must be stored in a refrigerator or cooling chamber at temperatures from 2°C to 8°C and out of light over the storage period.

Blisters with magnetic rods can be stored at 2°C to 25°C over the storage period. Shelf-life of the kit following the first opening of the primary container:

- the components of the kit must be stored in a refrigerator or a cooling chamber at temperatures from 2°C to 8°C over the storage period;

- blisters with magnetic rods can be stored at 2°C to 25°C over the storage period;

- cartridges, lysis solution, binding solution and wash solution No.1 must be stored in a refrigerator

or cooling chamber at temperatures from 2°C to 8°C and out of light over the storage period.

The kit stored under undue regime must not be used.

An expired the PREP-MB MAX DNA Extraction Kit must not be used.

We strongly recommend to follow the given instructions in order to obtain accurate and reliable results.

The conformity of the PREP-MB MAX DNA Extraction Kit to the prescribed technical requirements is subject to compliance of storage, transportation and handling conditions recommended by manufacturer.

| IV. TEST PROCEDURE AND QUALITY CONTROL |  |  |
| --- | --- | --- |
| 17. | Article 20.4.1(d) | Sample Type |

The PREP-MB MAX DNA Extraction Kit is designed to extract DNA from whole peripheral blood; smears/scrapings from urogenital tract and rectum; urine; ejaculate; milk; faeces.

General recommendations

- Use DNAse and RNAse free filter tips;

- When adding a solution to a sample containing biological material, be careful not to touch the walls of the tubes. If you touch the wall of the tube, change the tip. The tip should be changed each time the solution is removed from the sample.

- When transferring biological material into the extraction tube, rest the tip on the bottom of the tube. Avoid dripping from the tip onto the walls and cap of the tube.

- Open only the cap of the tube which is in work, then close the tube before proceeding to the next tube to prevent contamination.

Interfering substances

Maximum concentrations of interfering substances in biomaterial samples that do not affect the PCR assay: chlorhexidine (0.05% aqueous solution) – 10% v/v, high viscosity ultrasound gel – 5% v/v.

Sample collection

Smears/scrapes of epithelial cells

Sample taking is made with special sterile single-use tools – probes, cytobrushes, and swabs depending on the source of biological material according to established procedure.

ATTENTION! Remove mucus with sterile cotton swab before taking scrape from cervical channel. The taking of the sample is carried out:

- in tubes with transport medium intended by the manufacturer for transportation and storage of samples for PCR;

- in 1.5 mL tubes with 500 μL of a sterile physiological saline solution. Order of taking:

1. Open the tube.

2. Scrape epithelial cells from the corresponding biotope (urogenital tracts, rectum) with a sterile swab.

3. Put the swab into the tube with solution and rinse it thoroughly. Avoid spraying of solution.

4. Remove swab from solution, press it to the wall of tube and squeeze the rest of the liquid. Throw out the swab.

5. Close the tube tightly and mark it.

Blood sampling

Peripheral blood sampling is carried out in vacuum plastic tube. It may be 2.0 or 4.0 mL Vacuette blood collection tubes with anticoagulant, for example salt of EDTA at a final concentration of 2.0 mg/mL or sodium citrate anticoagulant. After taking the material, it is necessary to mix the blood with anticoagulant turning the tube 2 – 3 times.

ATTENTION! It is not allowed to use heparin as an anticoagulant.

The sample is ready for DNA extraction.

Ejaculate sampling

The ejaculate is obtained by masturbation and collected in a sterile container. The container with ejaculate is hermetically closed and marked.

Urine sampling

The portion of urine in the amount of 20–30 mL is selected for the analysis. The urine is taken into a special dry sterile container with volume of up to 60 mL, equipped with a hermetic screw-cap.

After the urine collection, container is tightly screwed and marked.

Milk sampling

- Collect the sample into the sterile container and close it tightly.

- Mix thoroughly and put 1.0 mL of the sample into the 1.5 mL tube.

- Centrifuge the tube at RCF(g) 12000-16000 for 3 minutes.

- Remove the supernatant, leaving approximately 100 μL in the tube.

- Close the tube tightly and mark it.

The sample is ready for DNA extraction.

Faeces sampling

Samples of faeces with mass (volume) 1.0-3.0 g (1.0-3.0 mL) are transferred to a special sterile dry flask by a single-use filtered pipette tip or single-use shovel. After sample collection the flask is tightly closed and marked.

Transportation and storage of the samples

Samples may be transported and stored at temperatures from 2 °C to 8 °C for no more than 24 hours prior to analysis. When it is impossible to deliver the material in the laboratory during the day, a one-time freezing of the material is allowed. Frozen material is allowed to be stored at temperatures from minus 18

°C to minus 22 °C for up to 3 months.

In case of usage transport media biological material samples are transported and stored according to the instruction for the transport medium used intended for subsequent sample analysis by PCR.

NOTE. If samples were stored at minus 18 °C to minus 22 °C, thaw them at room temperature (from 18 °C to 25 °C) or at 2 °C to 8 °C before DNA extraction.

Sample preparation Smears/scrapes of epithelial cells

- Centrifuge the tube at RCF(g) 16,000 for 10 minutes at room temperature (from 18°С to 25°С).

- Remove supernatant, leaving approximately 100 μL (precipitate+liquid fraction) in the tube.

Sample is ready for DNA extraction.

Ejaculate

Transfer 100 μL of liquid sample into a 1.5 mL plastic tube.

Sample is ready for DNA extraction.

Urine

- Transfer 1.0 mL of sample to 1.5 mL tube.

- Centrifuge the tube at RCF(g) 16000 for 10 minutes.

- Remove supernatant, leaving approximately 100 μL (precipitate+liquid fraction) in the tube. Samples are ready for DNA extraction.

Faeces

- Transfer approximately 50-100 mg (μL) of faeces into 1.5 mL tube with 1.0 mL of sterile physiological saline solution.

- Mix samples by vortexing for 10-15 seconds.

- Spin the tubes with samples for 30 seconds in vortex mixer.

- Transfer 100 μL of supernatant into the 1.5 mL tube. Samples are ready for DNA extraction.

To examine samples for the presence of bacteria – acute intestinal infections pathogens, further DNA extraction shall be carried out according to the PREP-MB MAX kit user instruction.

To examine samples for the presence of RNA-viruses – acute intestinal infections pathogens, the PREP-MB MAX kit IS NOT APPLICABLE!

To extract DNA of Gram-negative bacteria or to conduct a study of the intestinal biocenosis, it is necessary to pretreat the sample (faeces) with lysozyme solution:

- add 20 µL of 100 mg/mL lysozyme solution to each tube with biomaterial (faeces), vortex for 5-10 seconds and centrifuge for 1-3 seconds in vortex mixer;

- leave the tubes at room temperature between 18 °C and 25 °C for one hour;

NOTE. If necessary, samples with lysozyme can be stored at 2°C to 8°C for no more than 14 hours.

- extract DNA using the PREP-MB MAX kit.

| IV. TEST PROCEDURE AND QUALITY CONTROL |  |  |
| --- | --- | --- |
| 18. | Article 20.3(j) | Refer to Instructions for Use |
| 19. | Article 20.4.1(l) | Procedure Steps |

General requirements

1. Sorbent contains very light paramagnetic particles. It is recommended to use 1.0 mL tips to remove solutions from the samples using a single-channel pipette. When using a laboratory aspirator, set the aspiration rate to minimum and do not use large cross-sectional tips.

2. For sample preparation and extraction use DNAse and RNAse free filter tips.

3. Change the tip each time while removing, transferring or adding liquid into the tube. Do not touch the tube walls while adding liquid in the tube. If touching the wall, change the tip.

4. Always open the tube that you are working with, and close it after handling. It is not allowed to work simultaneously with several tubes with open caps.

5. Simultaneously with the extraction of DNA, a negative control should go through all stages of DNA extraction.

6. For information on the necessity of using an internal control, refer to the relevant user manual for PCR kits for the detection of pathogen DNA.

Assay procedure

1. A magnetic homogenizer is used.

2. Precipitate is allowed in the lysis solution, binding solution and wash solution № 1 when stored in the refrigerator (from 2 °C to 8 °C).

Prior to work:

  - check the lysis solution, binding solution, and wash solution № 1 for the absence of precipitate. If precipitate occurs, place the vial(s) on a thermostat preheated to 65 °C and warm until the precipitate is dissolved. Cool solutions to room temperature (from 18 °C to 25 °C) before use. The precipitate can also be dissolved at room temperature (from 18 °C to 25 °C) for approximately 12 hours;

  - stir the contents of the vial(s) by turning them upside down 5-10 times, avoiding foaming.

ATTENTION! When dispensing, use individual tips for each sample. Avoid sorbent loss during supernatant withdrawal.

1. Mark one 1.5 mL tube for each test sample and for the negative control “C-”.

NOTE. For pretreated samples with precipitate and supernatant obtained (epithelial cell scrapings, urine, etc.), mark test tubes with 100 µL of material prepared for assay.

2. Carefully remove protective foil from the required number of blister wells with magnetic rods.

3. Add a magnetic rod into each test tube using a forceps without touching the edge of the test tube.

4. When using internal control, add 10 µL of the corresponding internal control to each tube.

5. Stir all solutions thoroughly by inverting at least 20 times.

6. Add 150 µL of lysis solution to each tube without touching the edge of the tube.

ATTENTION! The solution should be applied to the bottom of the test tube, avoiding the formation of drops. If splashing occurs, it is necessary to spin down the drops on a vortex mixer.

7. Add 100 µL of the test sample into each test tube.

8. Add 100 µL of negative control from the kit to the tube labeled “C-”.

9. Close the caps tightly and place the tubes in the magnetic homogenizer. Start the homogenizer in pulse mode for 10 minutes.

ATTENTION! Each test tube should be positioned relative to the magnet so that the rod is always covered with solution, and uniform homogenization of suspension throughout the volume of liquid is ensured.

10. Add 250 µL of binding solution into each tube. Start the homogenizer in pulse mode for

- econds.

11. Carefully resuspend the sorbent in a vortex mixer. Turn the tube upside down and make sure that the sorbent does not stick to the bottom of the tube. If necessary, repeat stirring.

12. Add 60 µL of sorbent into each tube.

ATTENTION! Add the sorbent into each tube with a separate tip.

13. Place the tubes in the magnetic homogenizer. Start the homogenizer in pulse mode for 5 minutes.

ATTENTION! If the magnetic particles (sorbent) have partially deposited above and/or below the magnetic rod on the wall of the test tube during mixing in pulse mode, shake these test tubes for 3-5 seconds and centrifuge for 3-5 seconds in a vortex mixer, then return the test tubes to the magnetic homogenizer.

14. Without removing the tubes from the homogenizer, precipitate the sorbent on the walls of the tubes for 3 minutes (homogenizer off).

15. Collect as much as possible of the supernatant without removing the tubes from the magnetic homogenizer (with a separate tip from each tube). Avoid taking the sorbent.

16. Add 250 µL of wash solution № 1 to the precipitate. Close the caps of the tubes. Start the homogenizer in pulse mode for 30 seconds.

17. Without removing the tubes from the homogenizer, precipitate the sorbent onto the walls of the tubes for 3 minutes (homogenizer off).

18. Collect the supernatant without removing the tubes from the magnetic homogenizer (with a separate tip from each tube). Avoid taking the sorbent.

19. Repeat 16 – 18.

20. Add 500 µL of wash solution № 2 to the precipitate. Start the homogenizer in pulse mode

for 30 seconds.

21. Without removing the tubes from the homogenizer, precipitate the sorbent onto the walls of the tubes for 3 minutes (homogenizer off).

22. Collect as much supernatant as possible without removing the tubes from the magnetic homogenizer (with a separate tip from each tube). Avoid taking the sorbent.

23. Add 50 to 300 µL of elution solution into each tube. Shake the tubes in a vortex mixer so that

all particles from the walls are transferred to the solution.

The volume of elution solution depends on:

- type of assay;

- number of assay parameters;

- number of assays carried out from one DNA sample.

NOTE.The recommended amount of eluate for DNA extraction is given in the instructions for the PCR assays kits.

E x a m p l e :

| Type of assay | Amount of elution solution, µL |
| --- | --- |
| Virus DNA from whole blood | 50 |
| Integrated assays (e.g., Femobiome kits, Immunoquantex C/V, HPV QNT-21, genetic polymorphisms kits) | 300 |

24. Spin down the particles and the rod to the bottom of the tube in a vortex mixer.

25. Spin down the drops by centrifugation at RCF(g) 1,000-3,000 for 3-5 seconds.

26. Place the tubes in the magnetic homogenizer. Start the homogenizer in pulse mode for 15 minutes.

27. Without removing the tubes from the homogenizer, precipitate the sorbent onto the walls of the tubes for 3 minutes (homogenizer off).

Supernatant containing the extracted DNA is ready to be added to the reaction mixture for PCR amplification.

ATTENTION! When taking the supernatant containing DNA, do not remove the tubes from the magnetic homogenizer and do not touch the sorbent (magnetic particles).

The obtained DNA preparation can be stored at temperatures from 2°C to 8°C for up to 7 days or at temperatures from -18°C to -22°C for up to 1 year.

NOTE. If the DNA preparation is to be stored for more than 7 days, the supernatant must be transferred to a new tube.

The supernatant containing the extracted DNA is ready to be added to the reaction mixture for PCR amplification. The obtained DNA preparation can be stored at 2°C to 8°C for up to 7 days, or at -18°C to -22°C for up to one year.

NOTE.

- If the DNA preparation is to be stored for more than 7 days, the supernatant must be transferred to a new tube.

- If the cartridge is to be stored after running the DTstream program, it must be sealed with a protective lid and stored in accordance with the reagent kit storage conditions, avoiding the inversion of the cartridge.

| IV. TEST PROCEDURE AND QUALITY CONTROL |  |  |
| --- | --- | --- |
| 20. | Article 20.4.1(m) | Quality Control (QC) Procedure |

| V. RESULTS AND PERFORMANCE |  |  |
| --- | --- | --- |
| 21. | Article 20.4.1(m) | Interpretation of Results |

| V. RESULTS AND PERFORMANCE |  |  |
| --- | --- | --- |
| 22. | Article 20.4.1(q) | Analytical Performance Data |
| 23. | Article 20.4.1(q) | Clinical Performance Data |

a. The minimum amount of biomaterial for nucleic acids extraction is 100 µL. DNA yield from 100 µL sample: 5-50 ng.

b. Nucleic acid isolation purity: 1.6-2.0 OD260/280 nm.

c. Maximum sorbent capacity: at least 100 ng per extraction per DNA. Table 3 – Functional indicators by type of biomaterial.

| Type of biomaterial | Nucleic acid isolation purity OD260/280, nm | DNA yield from 100 µL sample |
| --- | --- | --- |
| Whole peripheral blood (n=45) | 1.6 -2.0 | 6-50 |
| Epithelial cell smears/scrapings from urogenital tract (n=40) | 1.6 -2.0 | 6-50 |
| Epithelial cell smears/scrapings from rectum (n=10) | 1.6 -2.0 | 7-40 |
| Urine (n=36) | 1.6 -2.0 | 8-49 |
| Ejaculate (n=27) | 1.6 -2.0 | 11-50 |
| Milk (n=5) | 1.6 -1.8 | 13-49 |
| Faeces (n=17) | 1.6 -2.0 | 8-42 |

d. Effectiveness of the reagent kit

The effectiveness of the PREP-MB MAX reagent kit was established during clinical trials with the use of additional amplification reagent kits (produced by “GENOMER”) in the biomaterial samples assay.

| Type of biomaterial | Number of test samples | Amplification reagent kit (abbreviated name) | The PREP-MB MAX reagent kit efficiency |
| --- | --- | --- | --- |
| Whole peripheral blood | 36 | CMV | 100% (91.65 - 100) |
|  | 43 | SIC* | 100% (92.92 - 100) |
| Epithelial cell smears/scrapings from urogenital tract | 41 | Mycoplasma hominis | 100% (92.60 - 100) |
| Epithelial cell smears/scrapings from rectum | 27 | Candida albicans | 100% (89.15 - 100) |
| Urine | 35 | CMV | 100% (91.40 - 100) |
| Ejaculate | 38 | Mycoplasma hominis | 100% (92.06 - 100) |
|  | 31 | SIC* | 100% (90.43 - 100) |
| Milk | 32 | CMV | 100% (90.70 - 100 |
| Faeces | 23 | Candida albicans | 100% (87.49 - 100) |
| The PREP-MB MAX reagent kit overall efficiency | 306 | 100% (98.94 – 100) |  |
| * The amount of extracted genomic DNA was sufficient for further PCR assays. |  |  |  |

| V. RESULTS AND PERFORMANCE |  |  |
| --- | --- | --- |
| 24. | Article 20.4.1(m) | Retesting Requirements |

| VI. ADDITIONAL INFORMATION |  |  |
| --- | --- | --- |
| 25. | Article 20.4.1(t) | Disposal of Used Materials |

| VI. ADDITIONAL INFORMATION |  |  |
| --- | --- | --- |
| 26. | Article 20.4.1(x) | Revision Date of Instructions for Use |

| VI. ADDITIONAL INFORMATION |  |  |
| --- | --- | --- |
| 27. | Article 20.4.1(v) | Standards with which Compliance is Claimed |

ISO 15223-1, ISO 18113 Reference to relevant standards such

| VI. ADDITIONAL INFORMATION |  |  |
| --- | --- | --- |
| 28. | Article 20.4.1(r) | Check Before Use |

Do not use the kit:

- When transportation and storage conditions are breached;

- When the reagents’ appearance does not respond to the kit passport;

- When kit components packaging is breached;

- After the expiry date provided.

Significant health effects are NOT anticipated from routine use of this kit when adhering to the instructions listed in the current manual.

| VI. ADDITIONAL INFORMATION |  |
| --- | --- |
| 29. | Technical Support Contact |

For technical support, please contact our dedicated technical support team by email: support@genomer.com.tr or Phone: +90 530 462 2637.

VII. SYMBOLS/PICTOGRAMS

|  | In vitro diagnostic medical device |  | Version |
| --- | --- | --- | --- |
|  | Temperature limit |  | Date of manufacture |
|  | Contains sufficient for<n>tests |  | Consult instructions for use |
|  | Use by date |  | Catalogue number |
|  | Batch code |  | Manufacturer |
|  | Non-sterile |  | Keep away from sunlight |
|  | Do not re-use |  | Caution |

# REVISION HISTORY

| Revision No. | Rev. Date | Description |
| --- | --- | --- |
| 00 | 14.11.2025 | First Issue |

# DOCUMENT APPROVAL

| Prepared By | Approved By |
| --- | --- |
| Quality Management Representative | General Manager |
| Elife ÖZDEMİR | Oktay ÖZTÜRK |