PRODUCT NAME:

PREP-MB-RAPID II DNA/RNA Extraction Kit

INTENDED USE:

The PREP-MB-RAPID II DNA/RNA Extraction Kit is designed to extract NA from biological materials: urine,

scrapes/smears of epithelial cells from urogenital tract, oropharynx, nasopharynx. This medical device is an auxiliary agent in clinical laboratory diagnostics.

Headquarters: Neighborhood of Dumlupınar, Nur St. No: 7/3 Post Code: 41250 Kartepe/Kocaeli/TURKEY

Branch: TUBITAK MAM Technology Free Zone, Neighborhood of Barıs SB, 5002 Street, New Tech. Building, Blok A, No: 4A/ Z09 Post Code: 41400

Gebze/Kocaeli/TURKEY

Trd Register Number: 34218

Email: info@genomer.com.tr

https://genomer.com.tr

Purpose

This document constitutes the Instructions for Use (IFU) for the PREP-MB-RAPID II DNA/RNA Extraction Kit, an In Vitro Diagnostic (IVD) medical device.

The European Union's In Vitro Diagnostic Medical Devices Regulation (IVDR) 2017/746 aims to ensure a high level of health and safety protection for patients and users throughout the entire lifecycle of the device.

These Instructions for Use have been prepared to fully satisfy all obligations stipulated under Annex I: General Safety and Performance Requirements of the IVDR, specifically within Chapter III: Information Supplied by the Manufacturer (Article 20).

The primary objective is to ensure the safe, accurate, and effective use of the device for its intended purpose. These instructions encompass all necessary technical, procedural, and legal information for laboratory professionals, including device performance, limitations, potential risks, and appropriate sample management.

Regulatory Compliance and Transparency

This guide not only employs pictograms and symbols strictly in accordance with national and international standards (specifically IVDR, EN ISO 15223-1, and EN ISO 18113) but also transparently presents all legally required device identifiers (UDI, Manufacturer, Authorized Representative).

Warning: It is imperative to read, understand, and meticulously adhere to these instructions in their entirety prior to using the device to ensure safe and accurate results. Failure to comply may lead to erroneous results, potential hazards, or invalidation of the CE marking.

| I. DESCRIPTIVE INFORMATION |  |  |  |
| --- | --- | --- | --- |
| 1. | Device Name, Trade Name, and Model | Article 20.3(a), 20.4.1(b) | PREP-MB-RAPID II DNA/RNA Extraction Kit |
| 2. | In Vitro Diagnostic Device | Article 20.3(b) | The PREP-MB-RAPID II DNA/RNA Extraction Kit is designed to extract NA from biological materials: urine, scrapes/smears of epithelial cells from urogenital tract, oropharynx, nasopharynx. This medical device is an auxiliary agent in clinical laboratory diagnostics. |
| 3. | Manufacturer Information | Article 20.3(d) | GENOMER BIOTECHNOLOGY R&D IND. TRD. LIMITED COMPANY, Headquarters: Neighborhood of Dumlupınar, Nur St. No: 7/3 Post Code: 41250 Kartepe/Kocaeli/TURKEY Branch: TUBITAK MAM Technology Free Zone, Neighborhood of Barıs SB, 5002 Street, New Tech. Building, Blok A, No: 4A/ Z09 Post Code: 41400 Gebze/Kocaeli/TURKEY Trd Register Number: 34218 Email: info@genomer.com.tr Web:https://genomer.com.tr |
| 4. | Authorized Representative Information | Article 20.3(e) | Required only if the manufacturer is located outside the EU. |
| 5. | Lot/Batch Number | Article 20.3(f) | UL- MBR-XX |
| 6. | Expiry/Use-by Date | Article 20.3(g) | YYYY/AA/GG |
| 7. | Catalog Number | Article 20.3(c) | I-MBR-v01-96 |
| 8. | Unique Device Identifier (UDI) Reference | Article 20.3(n) | 868512670IMBRMS |

The results obtained in accordance with the procedures and analytical acceptance criteria specified in the aforementioned Instructions for Use (IFU) guide confirm the conformity of the PREP-MB-RAPID II DNA/RNA Extraction Kit device with the General Safety and Performance Requirements (Annex I).

| II. INTENDED USE, SCOPE, AND PRINCIPLE |  |  |  |
| --- | --- | --- | --- |
| 9. | Intended Use and Function | Article 20.4.1(c) | INTENDED USE |

The PREP-MB-RAPID II DNA/RNA Extraction Kit is intended for semi-automatic and automatic extraction of human, bacterial, viral, and fungal NA for further analysis by PCR/ RT-PCR.

The PREP-MB-RAPID II DNA/RNA Extraction Kit is designed to extract NA from biological materials: urine, scrapes/smears of epithelial cells from urogenital tract, oropharynx, nasopharynx. This medical device is an auxiliary agent in clinical laboratory diagnostics.

Application of the kit does not depend on population and demographic aspects. There are no contradictions for use of the PREP-MB-RAPID II DNA/RNA Extraction Kit.

The PREP-MB-RAPID II DNA/RNA Extraction Kit can be used in clinical and diagnostic laboratories of medical institutions and research practice.

Potential users: personnel qualified in molecular diagnostics methods and working in clinical and diagnostic laboratory.

It is necessary to apply the kit only as directed in this instruction for use.

| II. INTENDED USE, SCOPE, AND PRINCIPLE |  |  |  |
| --- | --- | --- | --- |
| 10. | Test Principle | Article 20.4.1(e) | METHOD |

The method is based on lysis and release of nucleic acids under the action of a chaotropic agent (guanidine thiocyanate) with subsequent sorption on paramagnetic nanoparticles and washing to get rid of impurities.

The PREP-MB-RAPID II DNA/RNA Extraction Kit can be used in conjunction with medical devices designed for the PCR/RT-PCR analysis of nucleic acids. It is recommended to validate the PREP-MB-RAPID II DNA/RNA Extraction Kit with the reagent kit for PCR/RT-PCR analysis of nucleic acids.

| II. INTENDED USE, SCOPE, AND PRINCIPLE |  |  |  |
| --- | --- | --- | --- |
| 11. | Adequate Number of Tests | Article 20.3(k) | CONTENTS |

The PREP-MB-RAPID II DNA/RNA Extraction Kit content is represented in Table 1.

| Reagent | Description | Total volume | Amount |
| --- | --- | --- | --- |
| Lysis solution | Blue transparent foamy liquid | 28.8 mL (300 μL in each well) | 1*96 deep-well plate |
| Wash solution | Green transparent liquid | 57.6 mL (600 μL in each well) | 1*96 deep-well plate |
| Elution solution | Colorless or pink transparent liquid | 9.6 mL (100 μL in each well) | 1*96 deep-well plate |
| Sorbent | Liquid with precipitate forming brown suspension upon shaking | 4.8 mL (1.2 mL in each tube) | 4 tubes |
| Elution solution | Colorless or pink transparent liquid | 19.2 mL | 1 vial |
| 96 tip comb | 1 pc |  |  |
| 96 deep-well plate* | 1 pc |  |  |
| PCR seal sheet** | 1 pc |  |  |
| * for magnetic rod tips discharge ** to seal plates with extracted nucleic acids during storage |  |  |  |

Reagent kit is single-use and is designed for NA extraction from 96 test samples (one run of 96 samples), including controls.

REAGENTS AND EQUIPMENT REQUIRED BUT NOT PROVIDED

Specimen collection

- Sterile single use swabs, cotton swabs e.t.c for sampling of biomaterial;

- Sterile tubes containing transport media: “GENOMER” made STOR-F or equivalent for the transportation of the sample.

NA extraction

- Biological safety cabinet class II;

- Refrigerator with freezer;

- Vortex mixer;

- High speed centrifuge (RCF(g) at least 12,000) for 1.5 mL tubes;

- Laboratory aspirator with trap flask for the removal of supernatant;

- RNase and DNase free non-filtered pipette tips for aspirator with trap flask;

- Solid-state thermostat (temperature from 65 °C);

- Single channel pipettes (dispensers covering 20-1,000 μL volume range);

- RNase and DNase free filtered pipette tips (volume 200 μL; 1,000 μL);

- Pipette stand;

- Magnetic rack;

- Magnetic pick-up tool or plastic pick-up tool for magnetic rods;

- 1.5 mL RNase and DNase free micro-centrifuge tubes (for DTstream micro-centrifuge tubes SSI-1260 are recommended);

- Tube rack for 1.5 mL tubes;

- Dosing  instrument  DTstream  12L4  or  15L41  configuration  at  least  1.1  (for  packages  A P-122-A/9EU, P

- RNase and DNase free filtered pipette tips (volume 1000 μL) for DTstream (for packages A

P-122-A/9EU, P);

- System for automatic nucleic acid extraction in 96 deep-well plate (Allsheng Auto-Pure 96, KingFisher Flex);

- Adhesive seal sheet (for example, 4titude 4ti-0500) (except package P, No. 2  Р-124-P/9EU);

- 96 tip comb;

- Physiological saline solution 0.9% NaCl sterile;

- Transport medium for transport and storage of PCR biomaterial (GENOMER STOR-F is recommended)

- Container for used pipette tips, tubes and other consumables;

- Container for used pipette tips for DTstream;

- Powder-free surgical gloves;

- Disinfectant solution.

| III. SAFETY AND PRECAUTIONS |  |  |  |
| --- | --- | --- | --- |
| 12. | Warnings and Precautions | Article 20.4.1(n) | All necessary general warnings for safe use |
| 13. | Biohazard Warning | Article 20.4.1(p) | Specimens pose a potential infection risk |
| 14. | Hazardous Chemical Substances | Article 20.4.1(o) | If kit components contain hazardous substances |
| 15. | Restriction on Use | Article 20.1(f) | For use only by trained laboratory personnel |

Only personnel trained in the methods of molecular diagnostics and the rules of work in the clinical and diagnostic laboratory are allowed to work with the kit.

Handle and dispose all biological samples, reagents and materials used to carry out the assay as if they were able to transmit infective agents. The samples must be exclusively employed for certain type of analysis. Samples must be handled under a laminar flow hood. Tubes containing different samples must never be opened at the same time. Pipettes used to handle samples must be exclusively employed for this specific purpose. The pipettes must be of the positive dispensation type or be used with aerosol filter tips. The tips employed must be sterile, free from the DNases and RNases, free from DNA and RNA. The reagents must be handled under a laminar flow hood. The reagents required for amplification must be prepared in such a way that they can be used in a single session. Pipettes used to handle reagents must be exclusively employed for this specific purpose. The pipettes must be of the positive dispensation type or be used with aerosol filter tips. The tips employed must be sterile, free from the DNases and RNases, free from DNA and RNA. Avoid direct contact with the biological samples reagents and materials used to carry out the assay. Wear powder-free surgical gloves. Wear protective clothing (work clothes and personal protective equipment) working with microorganisms classified as particularly pathogenic. The protective clothing and personal protective equipment must comply with the work to be performed and health and safety requirements. Avoid producing spills or aerosol. Any material being exposed to biological samples must be treated for at least 30 minutes with disinfecting solution or autoclaved for 1 hour at 121 °C before disposal.

Molecular biology procedures, such as nucleic acids extraction, reverse transcription, PCR-amplification and detection require qualified staff to avoid the risk of erroneous results, especially due to the degradation of nucleic acids contained in the samples or sample contamination by amplification products.

All the liquid solutions are designed for single use and can not be used more than once in amplification reactions. Plastic tubes do not contain phthalates. Do not breathe gas/fumes/vapor/spray produced by the components of the kit. Do not eat/drink components of the kit. Avoid contact with eyes. Only use the reagents provided in the kit and those recommended by manufacturer. Do not mix reagents from different batches. Do not use reagents from third party manufacturers’ kits. All laboratory equipment, including pipettes, test tube racks, laboratory glassware, lab coats, bouffant caps, etc., as well as reagents should be strictly stationary. It is not allowed to move them from one room to another. Equip separate areas for the extraction/preparation of amplification reactions and for the amplification/detection of amplification products. Never introduce an amplification product in the area designed for extraction/preparation of amplification reactions. Wear lab coats, gloves and tools, which are exclusively employed for the extraction/preparation of the amplification reaction and for the amplification/detection of the amplification products. Never transfer lab coats, gloves and tools from the area designed for amplification/detection of the amplification products to the area designed for extraction/preparation of amplification reactions.

Emergency actions

Eye Contact: If any component of this kit enters the eyes, wash eyes gently under potable running water for 15 minutes or longer, making sure that the eyelids are held open. If pain or irritation occurs, obtain medical attention.

Skin Contact: If any component of this kit contacts the skin and causes discomfort, remove any contaminated clothing. Wash affected area with plenty of soap and water. If pain or irritation occurs, obtain medical attention.

Ingestion: If any component of this kit is ingested, wash mouth out with water. If irritation or discomfort occurs, obtain medical attention.

| III. WARNINGS AND PRECAUTIONS |  |  |  |
| --- | --- | --- | --- |
| 16. | Special Storage Conditions | Article 20.4.1(j) | Required temperature range for reagent storage |

Expiry date – 12 months from the date of production.

The PREP-MB-RAPID II DNA/RNA Extraction Kit must be transported in thermoboxes with ice packs by all types of roofed transport at temperatures inside the thermoboxes corresponding to storage conditions of the kit components.

The kit must be transported in the upright position in accordance with the handling sign “TOP”. Kits transported with violation of temperature conditions must not be used.

All components of the PREP-MB-RAPID II DNA/RNA Extraction Kit must be stored at temperatures from

2 °C to 25 °C and out of light over the storage period.

When stored in refrigerator (from 2 °С to 8 °С), a minor precipitate is allowed in lysis solution.

The excessive temperature and light can be detrimental to product performance.

Shelf-life of the kit following the first opening of the primary container: the components of the kit must be stored at temperatures from 2 °C to 25 °C and out of light over the storage period.

The kit must be stored in the upright position in accordance with the handling sign “TOP”. The kit stored under undue regime shall not be used.

An expired PREP-MB-RAPID II DNA/RNA Extraction Kit must not be used.

We strongly recommend to follow the given instructions in order to obtain accurate and reliable results.

Conformity of the PREP-MB-RAPID II DNA/RNA Extraction Kit to the prescribed technical requirements is subject to compliance of storage, transportation and handling conditions recommended by manufacturer.

| IV. TEST PROCEDURE AND QUALITY CONTROL |  |  |  |
| --- | --- | --- | --- |
| 17. | Sample Type | Article 20.4.1(d) | Type of specimen accepted for testing |

The PREP-MB-RAPID II DNA/RNA Extraction Kit is designed to extract DNA/RNA from urine, scrapes/smears of epithelial cells from urogenital tract, oropharynx, nasopharynx.

## General requirements

- During biomaterial preparation and NA extraction use DNAse and RNAse free tips (filter tips, except for supernatant collection using aspirator);

- When adding solution to the tube containing biomaterial, introduce it carefully, without touching the walls of the tube. If touching occurred, change the tip. Change the tip after each solution removal from the sample.

- Only open the cap of the tube you are working with, then close the tube before proceeding to the next tube to prevent contamination.

Interfering substances

Concentrations of interfering compounds that do not influence subsequent reverse transcription and polymerase chain reaction:

Whole blood 5.0% v/v, mucus (mucin), chlorhexidine (water solution 0.05%), Lasolvan Rhino (nasal spray), Rhinofluimucil (nasal spray), Tysine (nasal spray), Oxoline (nasal ointment), Pinosol (nasal spray), Tantum Verde (spray), Hexoral (aerosol), Berodual (aerosol), Salbutamol-Teva (aerosol), Grippferon (nasal spray) – 10% v/v.

| Biomaterial | Interfering substance | Examined concentration in the sample |
| --- | --- | --- |
| Endogenous substances |  |  |
| Urine, smears/scrapes of epithelial cells from urogenital tract, oropharynx, nasopharynx | Hemoglobin | 0.35 mg/ml |
| Smears/scrapes of epithelial cells from urogenital tract, oropharynx, nasopharynx | Mucus (mucin) | 10% v/v |
| Exogenous substances |  |  |
| Urine, smears/scrapes of epithelial cells oropharynx, nasopharynx | Chlorhexidine (water solution 0.05%) | 10% v/v |
| Smears/scrapes of epithelial cells oropharynx, nasopharynx | Lasolvan Rhino (nasal spray) | 10% v/v |
| Smears/scrapes of epithelial cells oropharynx, nasopharynx | Rhinofluimucil | 10% v/v |
| Smears/scrapes of epithelial cells oropharynx, nasopharynx | Tysine (nasal spray) | 10% v/v |
| Smears/scrapes of epithelial cells oropharynx, nasopharynx | Oxoline (nasal ointment) | 10% v/v |
| Smears/scrapes of epithelial cells oropharynx, nasopharynx | Pinosol (nasal spray) | 10% v/v |
| Smears/scrapes of epithelial cells oropharynx, nasopharynx | Tantum Verde (spray) | 10% v/v |
| Smears/scrapes of epithelial cells oropharynx, nasopharynx | Hexoral (aerosol) | 10% v/v |
| Smears/scrapes of epithelial cells oropharynx, nasopharynx | Berodual (aerosol) | 10% v/v |
| Smears/scrapes of epithelial cells oropharynx, nasopharynx | Salbutamol-Teva (aerosol) | 10% v/v |
| Smears/scrapes of epithelial cells oropharynx, nasopharynx | Grippferon (nasal spray) | 10% v/v |

Method limitations: local use of medications (sprays, drops, creams and ointments), vaginal ultrasound less than 24 hours before the assay. When using aerosols and other forms of drugs for inhalation in the treatment of bronchial asthma, material for assay should be taken at least three hours after inhalation.

Sample collection

WARNING! Before NA extraction pre-processing of biomaterial samples is required.

Urine

The first portion of morning urine in the amount of 20–30 mL is collected for the analysis. The urine is taken into a special dry sterile container with volume of up to 60 mL, equipped with a hermetical screw- cap.

After urine collection, container is tightly screwed and marked.

Scrapes/smears of epithelial cells from urogenital tract

Sample intake is made with special sterile single-use tools (e.g., probes).

Method limitations: Local application of medications, use of lubricants, vaginal ultrasound less than 24 hours before the assay.

Sample intake is carried out in accordance with the instruction to the used transport medium.

## Taking urogenital scrapes

On the eve of the examination, women should not perform vaginal douching or sprays. In order to obtain an accurate result, it is necessary that biomaterial contains as many epithelial cells as possible and a minimum amount of mucus and blood. Incorrect biomaterial sampling may lead to an unreliable result and, thus, to the need to repeat the biomaterial collection.

WARNING! Before obtaining a scrape of epithelial cells from the urethra, posterolateral vaginal vault and cervical canal, remove free-flowing secretion with a sterile cotton swab.

## Vaginal sampling

The sample must be taken prior to manual examination. Speculum can be treated with hot water before examination. Antiseptics must not be used for speculum treatment. The sample must be taken from the posterolateral vaginal vault. For girls, the material is taken from the mucous membrane of the vaginal vestibule, and in some cases from the posterior vaginal vault through the hymenal rings.

## Urethral sampling

Patient must not urinate within 1.5-2 hours prior to sampling procedure. The external urethral orifice must be treated with a swab moistened with sterile physiological saline solution just prior to the sampling procedure.

In the case of purulent discharge, the sample must be taken 15-20 minutes after urinating. In the absence of discharge, it is necessary to massage urethra with sampling swab or brush. Carefully insert the swab into the woman’s urethra to a depth of 1-1.5 cm. A child’s sample must be taken from the external urethral orifice.

## Cervical sampling

Remove mucus with a swab prior to sampling, and treat the cervix with sterile physiological saline solution. Carefully insert sampling swab into the cervix to a depth of 0.5-1.5 cm. Avoid contact with vaginal wall when removing the swab.

## Oropharyngeal smears

Take the smear with a swab with a rotational movement from the surface of the tonsils, palatine arches and the back wall of the pharynx.

## Nasopharyngeal smears

Smears/scrapes are taken with a dry sterile probe, the probe is inserted with a slight movement along the outer wall of the nose to a depth of 2-3 cm to the lower nasal shell. Then the probe is slightly lowered to the bottom, inserted into the lower nasal passage under the lower nasal shell, and removed along the outer nasal wall after a rotating movement.

Transport and storage of samples

Urine

Urine samples must be transported and stored:

- at temperature from 2 °С to 8 °С for no longer than 1 day.

- at temperature from minus 18 °С to minus 22 °С for no longer than one week.

WARNING! Only one freezing-thawing of material is allowed.

Scrapes/smears of epithelial cells from urogenital tract

Transport and storage conditions for urogenital scrapes are determined by instructions for use for the transport media used for transport and storage of samples.

## Oropharyngeal, nasopharyngeal smears

| Type of the sample | Collecting material requirements | Transportation | Storage conditions before transportation | Сomments |
| --- | --- | --- | --- | --- |
| Nasopharyngeal and oropharyngeal smears | Plastic test tubes and swab for smears** | 4 °C | ≤5 days: 4 °C >5 days*: minus 70 °C | Nasopharyngeal and oropharyngeal swabs should be placed in the same tube to increase the viral load |

* - If it is not possible to store samples at minus 70 °C, store samples at minus 20 °C.

** - Use a transport medium for storage and transportation of respiratory smears, or saline solution (if the sample is transported to the laboratory no more than 24 hours after sample taking), or a dry swab (if the sample is transported to the laboratory no more than 4 hours after sample taking).

WARNING! Only one freezing-thawing unfreezing of the material is allowed.

Sample preparation

Biomaterial preparation (if necessary) is performed in accordance with instructions for use for the transport media used for transport and storage of samples.

Preparation of epithelial smears/scrapes in STOR-F transport media, PreservCyt transport media (solution) (Hologic Inc, USA), Cellprep LBC Solution (Biodyne, South Korea) (if necessary), and urine is performed in accordance with Table 5.

| Biomaterial | NA extraction using reagent kit |  |
| --- | --- | --- |
|  | Package N Package A Package P (biomaterial introduced manually into plate No. 1) | Package P (biomaterial introduced automatically into plate No. 1) |
| Scrapes of epithelial cells from urogenital tract in STOR-F transport medium | Centrifuge the tube with biomaterial at RCF(g) 12,000-16,000 for 3 minutes at room temperature (from 18 °С to 25 °С), then remove supernatant leaving ~100 μL (precipitate + liquid fraction) | Biomaterial centrifugation in STOR-F transport medium before NA extraction is not required |
| Scrapes of epithelial cells from urogenital tract in PreservCyt transport medium (solution) (Hologic Inc, USA), Cellprep LBC Solution (Biodyne, South Korea) | - Mix thoroughly the sample in the transport medium (solution) by shaking the vial intensively. - Transfer 1.0 mL of biomaterial into 1.5 mL plastic tube. Close the cap tightly. - Centrifuge the tube at RCF(g) 12,000-16,000 for 3 minutes. - Remove supernatant as fully as possible. |  |
|  | 5. Add 100 μL of STOR-F transport medium or sterile physiological saline solution to precipitate and resuspend the precipitate by pipetting. | 5. Add 200 μL of STOR-F transport medium or sterile physiological saline solution to precipitate and resuspend the precipitate by pipetting. |
| Urine | - Transfer 1.0 mL of urine into 1.5 mL plastic tube. Close the tube tightly. - Centrifuge the tube at RCF(g) 12,000-16,000 for 3 minutes. - If after centrifugation the volume of salts precipitate exceeds 100 μL (more than 1/3 of the tube), remove supernatant as fully as possible and do 4-6. If precipitate volume does not exceed 100 μL, remove supernatant leaving up to 100 μL in the tube (precipitate + liquid fraction). The sample is ready for NA extraction. - Add 1.0 mL of sterile physiological saline solution to precipitate. Close the tube tightly. - Centrifuge the tube at RCF(g) 12,000-16,000 for 3 minutes. |  |

|  | 6. Remove supernatant leaving up to 100 μL in the tube (precipitate + liquid fraction). | 6. Remove supernatant leaving up to 100 μL in the tube (precipitate + liquid fraction). 7. Add 100 μL of STOR-F transport medium or sterile physiological saline solution to precipitate and resuspend the precipitate by pipetting. |
| --- | --- | --- |
| Smears of epithelial | Biomaterial centrifugation in STOR-F transport medium before NA extraction is not |  |
| cells from | required |  |
| oropharynx, |  |  |
| nasopharynx in |  |  |
| STOR-F transport |  |  |
| medium |  |  |

| IV. TEST PROCEDURE AND QUALITY CONTROL |  |  |  |
| --- | --- | --- | --- |
| 18. | Refer to Instructions for Use | Article 20.3(j) | Must refer to instructions for detailed procedure |
| 19. | Procedure Steps | Article 20.4.1(l) | Flowchart from nucleic acid extraction to qPCR cycling |

WARNING!

1. To introduce and add reagents and samples use RNase and DNase free filter tips.

2. Change the tips after each solution removal from the tube.

3. To prevent contamination only open the tube that you are working with (adding sample/reagent, removing supernatant), and close it after handling. It is not allowed to work simultaneously with several tubes with open caps.

4. Add solution to the tube/well containing biomaterial carefully, without touching the walls of the tube/well. If touching occurred, change the tip.

5. If samples are introduced automatically, make sure that tubes with samples do not contain probe residues.

6. Treat test samples and controls the same way and simultaneously according to this instruction.

7. Lysis solution can form precipitate if stored in the refrigerator (from 2°С to 8°С). Dissolve it by placing vial, cartridge or plate onto a thermostat warmed to 65 °С and warm until full precipitate dissolution. Cool the solution to room temperature (from 18°С to 25°С) before work. Alternatively, the precipitate can be dissolved at room temperature (from 18°С to 25°С) in the course of 12 hours.

8. Cap opening may occur during tube heating! Use tubes with locking caps (Eppendorf Safe-Lock Tubes) or programmable thermostat with clamp cover (solid-state programmable thermostat TT-1-DNA-Tech manufactured by DNA-Technology R&P, LLC).

Assay procedure for package P (sets No. 1 and 2)

1. Mark 3 96 deep-well plate from the reagent kit:

      - No. 1 with lysis solution

      - No. 2 with wash solution

      - No. 3 with elution solution

2. In case of using the PREP-MB-RAPID II with reagent kit for acute viral respiratory infection agents’ nucleic acids detection by RT-PCR (“DNA-Technology”, LLC) that includes internal control RNA-IC “A”, add 250 μL of internal control into each tube with sorbent, close the tubes and vortex for 3-5 seconds.

3. Spin down the drops and condensate from the walls of deep-well plates on vortex at RCF(g) 800 for 1-3 minutes, if necessary.

Note. Use dee-well plates with distilled water sealed with adhesive seal as a counterbalance.

4. Remove the protective sealing from the plates.

5. Resuspend the sorbent by pipetting and add 60 μL of the mixture of the sorbent and internal control or 50 μL of the sorbent in each well of the plate No. 1.

6. Add 100 μL of test samples in the corresponding wells of the plate No.1 containing lysis solution and sorbent.

Note. For samples pretreated with obtaining precipitate, resuspend the sample by pipetting and add into the corresponding plate well.

7. Add 100 μL of transport medium (for example, STOR-F) or sterile physiological saline solution into the well for negative control (C-).

8. If positive control is intended for NA extraction, add 100 μL of positive control into the corresponding well.

9. Prepare the automatic NA extraction system by setting the prepared deep-well plates and the necessary consumables (tip comb, empty 96 deep-well plate for disposing of the tip comb) according to the operation manual.

10. Select and run the NA extraction script.

11. Wait until the system finishes operation.

NA preparation is ready for PCR/RT-PCR (it is in plate No. 3 with elution solution). Use adhesive plate seal for storage of the plate with the extracted NA.

If necessary add 200 µL of elution solution in each well of the plate No. 3 (see example below). When adding the solution into the well with NA, add liquid carefully, without touching the walls of the well. If touching occurred, change the tip.

The final volume of elution solution depends on:

- type of assay;

- number of assays carried out from one NA sample.

E x a m p l e :

| Type of assay | Amount of elution solution, µL |
| --- | --- |
| Human acute respiratory viral infections pathogens nucleic acids RT-PCR detection kits (“DNA-Technology”), real-time PCR DNA detection kits (“DNA-Technology”) | 100 (default volume of solution in deep-well plates, no adding required) |
| Multiplex assays (e.g., Femoflor®, Androflor®, HPV-QUANT-21® kits) | 300 |

| IV. TEST PROCEDURE AND QUALITY CONTROL |  |  |  |
| --- | --- | --- | --- |
| 20. | Quality Control (QC) Procedure | Article 20.4.1(m) | Details for use of Positive, Negative and Internal Control (IC) |

| V. RESULTS AND PERFORMANCE |  |  |  |
| --- | --- | --- | --- |
| 21. | Interpretation of Results | Article 20.4.1(m) | Ct cut-off values and result table for each pathogen |

| V. RESULTS AND PERFORMANCE |  |  |  |
| --- | --- | --- | --- |
| 22. | Analytical Performance Data | Article 20.4.1(q) | Technical parameters such as LOD, Specificity, Precision, Linear Range. |
| 23. | Clinical Performance Data | Article 20.4.1(q) | Diagnostic Sensitivity, Specificity, and expected values (prevalence) in pediatric population. |

- The minimum amount of biomaterial for nucleic acids extraction:

- Transport medium STOR-F containing smears/scrapes of epithelial cells from urogenital tract, oropharynx, nasopharynx: 100 μL;

- Transport medium (solution) PreservCyt (Hologic Inc, USA), Cellprep LBC Solution (Biodyne, South Korea) containing smears/scrapes of epithelial cells from urogenital tract: 1.0 mL;

- urine: 1.0 mL.

- Functional characteristics of the reagent kit:

- Purity of nucleic acid samples (А260/280) – 1.5-2.4;

- Concentration of nucleic acids in 100 μL of preparation is at 5.9-92.8 ng/μL of nucleic acid solution.

- Effectiveness of the reagent kit

The effectiveness of the reagent kit was confirmed during clinical trials using additional reagent kits in for study of biological material samples:

| Biomaterial | Test samples, pcs | Amplification kit (short name) | Effectiveness of PREP-MB-RAPID II |
| --- | --- | --- | --- |
| Scrapes/smears of epithelial cells from oropharynx, nasopharynx | 26 | AVRI Multiplex | 100 % (Рtrue = 91.52) |
|  | 32 | Candida albicans | 100 % (Рtrue = 93.06) |
|  | 32 | HSV | 100 % (Рtrue = 93.06) |
|  | 30 | BacScreen OM | 100 % (Рtrue = 92.62) |
| Scrapes/smears of epithelial cells from urogenital tract | 24 | Femoflor® Screen | 100 % (Рtrue = 90.85) |
|  | 23 | HPV-Quant | 100 % (Рtrue = 90.47) |
| urine | 30 | Androflor® Screen | 100 % (Рtrue = 92.62) |
|  | 32 | HSV1/HSV2/CMV Multiplex | 100 % (Рtrue = 93.06) |
|  | 32 | Candida albicans | 100 % (Рtrue = 93.06) |
| General effectiveness of PREP-MB-RAPID II | 261 | - | 100 % (Рtrue = 99.12) |

- Within-batch and between-batch precision

Within-batch precision – 100 % (88.43% - 100%).

Between-batch precision – 100 % (88.43-% - 100%)

| V. RESULTS AND PERFORMANS |  |  |  |
| --- | --- | --- | --- |
| 24. | Retesting Requirements | Article 20.4.1(m) | How to repeat invalid or equivocal results. |

| VI. ADDITIONAL INFORMATION |  |  |  |
| --- | --- | --- | --- |
| 25. | Disposal of Used Materials | Article 20.4.1(t) | Instructions for medical and hazardous waste management |

Remove waste materials (tubes, tips) only in a special closed container containing a disinfectant solution. Work surfaces, as well as rooms where NA extraction and PCR are performed, must be irradiated with bactericidal irradiators for 30 minutes before and after the work.

Waste materials are disposed of in accordance with local and national standards. All surfaces in the laboratory (work tables, test tube racks, equipment, etc.) must be treated daily with disinfecting solution.

| VI. ADDITIONAL INFORMATION |  |  |  |
| --- | --- | --- | --- |
| 26. | Revision Date of Instructions for Use | Article 20.4.1(x) | 2025-11-25 |
| 27. | Standards with which Compliance is Claimed | Article 20.4.1(v) | ISO 15223-1, ISO 18113 Reference to relevant standards such |

| VI. ADDITIONAL INFORMATION |  |  |  |
| --- | --- | --- | --- |
| 28. | Check Before Use | Article 20.4.1(r) | Actions in case of damaged packaging or reagent deterioration |

Do not use the kit:

- When transportation and storage conditions are breached;

- When the reagents’ appearance does not respond to the kit passport;

- When kit components packaging is breached;

- After the expiry date provided.

Significant health effects are NOT anticipated from routine use of this kit when adhering to the instructions listed in the current manual.

| VI. ADDITIONAL INFORMATION |  |  |  |
| --- | --- | --- | --- |
| 29. | Technical Support Contact | - | Contact details for questions and technical issues. |

For technical support, please contact our dedicated technical support team by email: support@genomer.com.tr or Phone: +90 530 462 2637.

VII. SYMBOLS/PICTOGRAMS

|  | In vitro diagnostic medical device |  | Version |
| --- | --- | --- | --- |
|  | Temperature limit |  | Date of manufacture |
|  | Contains sufficient for<n>tests |  | Consult instructions for use |
|  | Use by date |  | Catalogue number |
|  | Batch code |  | Manufacturer |
|  | Authorized representative in the European Community |  | Caution |
|  | Non-sterile |  | Do not re-use |

# REVISION HISTORY

| Revision No. | Rev. Date | Description |
| --- | --- | --- |
| 00 | 14.11.2025 | First Issue |

# DOCUMENT APPROVAL

| Prepared By | Approved By |
| --- | --- |
| Quality Management Representative | General Manager |
| Elife ÖZDEMİR | Oktay ÖZTÜRK |