PRODUCT NAME:

STOR F Transport Medium

INTENDED USE:

The STOR F Transport Medium is intended for transport and storage of human biological samples (scrapes/smears of epithelial cells from urogenital tract, oropharynx, nasopharynx, rectum, skin, conjunctiva of the eye) followed by nucleic acids analysis (human DNA, DNA of microorganisms, RNA of viruses) by polymerase chain reaction method.

Headquarters: Neighborhood of Dumlupınar, Nur St. No: 7/3 Post Code: 41250 Kartepe/Kocaeli/TURKEY

Branch: TUBITAK MAM Technology Free Zone, Neighborhood of Barıs SB, 5002 Street, New Tech. Building, Blok A, No: 4A/ Z09 Post Code: 41400

Gebze/Kocaeli/TURKEY

Trd Register Number: 34218

Email: info@genomer.com.tr

https://genomer.com.tr

Purpose

This document constitutes the Instructions for Use (IFU) for the STOR F Transport Medium, an In Vitro Diagnostic (IVD) medical device.

The European Union's In Vitro Diagnostic Medical Devices Regulation (IVDR) 2017/746 aims to ensure a high level of health and safety protection for patients and users throughout the entire lifecycle of the device.

These Instructions for Use have been prepared to fully satisfy all obligations stipulated under Annex I: General Safety and Performance Requirements of the IVDR, specifically within Chapter III: Information Supplied by the Manufacturer (Article 20).

The primary objective is to ensure the safe, accurate, and effective use of the device for its intended purpose. These instructions encompass all necessary technical, procedural, and legal information for laboratory professionals, including device performance, limitations, potential risks, and appropriate sample management.

Regulatory Compliance and Transparency

This guide not only employs pictograms and symbols strictly in accordance with national and international standards (specifically IVDR, EN ISO 15223-1, and EN ISO 18113) but also transparently presents all legally required device identifiers (UDI, Manufacturer, Authorized Representative).

Warning: It is imperative to read, understand, and meticulously adhere to these instructions in their entirety prior to using the device to ensure safe and accurate results. Failure to comply may lead to erroneous results, potential hazards, or invalidation of the CE marking.

| I. DESCRIPTIVE INFORMATION |  |  |  |
| --- | --- | --- | --- |
| 1. | Device Name, Trade Name, and Model | Article 20.3(a), 20.4.1(b) | STOR F Transport Medium |
| 2. | In Vitro Diagnostic Device | Article 20.3(b) | The STOR F Transport Medium is intended for transport and storage of human biological samples (scrapes/smears of epithelial cells from urogenital tract, oropharynx, nasopharynx, rectum, skin, conjunctiva of the eye) followed by nucleic acids analysis (human DNA, DNA of microorganisms, RNA of viruses) by polymerase chain reaction method. |
| 3. | Manufacturer Information | Article 20.3(d) | GENOMER BIOTECHNOLOGY R&D IND. TRD. LIMITED COMPANY, Headquarters: Neighborhood of Dumlupınar, Nur St. No: 7/3 Post Code: 41250 Kartepe/Kocaeli/TURKEY Branch: TUBITAK MAM Technology Free Zone, Neighborhood of Barıs SB, 5002 Street, New Tech. Building, Blok A, No: 4A/ Z09 Post Code: 41400 Gebze/Kocaeli/TURKEY Trd Register Number: 34218 Email: info@genomer.com.tr Web:https://genomer.com.tr |
| 4. | Authorized Representative Information | Article 20.3(e) | Required only if the manufacturer is located outside the EU. |
| 5. | Lot/Batch Number | Article 20.3(f) | UL- STRF-XX |
| 6. | Expiry/Use-by Date | Article 20.3(g) | YYYY/AA/GG |
| 7. | Catalog Number | Article 20.3(c) | S-STRF-v01-100 |
| 8. | Unique Device Identifier (UDI) Reference | Article 20.3(n) | 868512670SSTRFTT |

The results obtained in accordance with the procedures and analytical acceptance criteria specified in the aforementioned Instructions for Use (IFU) guide confirm the conformity of the STOR F Transport Medium device with the General Safety and Performance Requirements (Annex I).

| II. INTENDED USE, SCOPE, AND PRINCIPLE |  |  |  |
| --- | --- | --- | --- |
| 9. | Intended Use and Function | Article 20.4.1(c) | INTENDED USE |

The STOR F Transport Medium is intended for transport and storage of human biological samples (scrapes/smears of epithelial cells from urogenital tract, oropharynx, nasopharynx, rectum, skin, conjunctiva of the eye) followed by nucleic acids analysis (human DNA, DNA of microorganisms, RNA of viruses) by polymerase chain reaction method.

This medical device is an auxiliary agent in clinical laboratory diagnostics and is intended for transport and storage of human biological samples, followed by nucleic acids analysis by polymerase chain reaction method.

The application of the kit does not depend on population and demographic aspects. There are no contradictions for use of the STOR F Transport Medium.

The STOR F Transport Medium can be used in clinical and diagnostic laboratories of medical institutions and research practice.

Potential users: personnel qualified in molecular diagnostics methods and working in the clinical and diagnostic laboratory.

It is necessary to apply the kit only as directed in this instruction for use.

| II. INTENDED USE, SCOPE, AND PRINCIPLE |  |  |  |
| --- | --- | --- | --- |
| 10. | Test Principle | Article 20.4.1(e) | METHOD |

The STOR F Transport Medium is a ready-to-use isotonic water-salt solution with the addition of a preservative, which prevents the growth of microorganisms.

| II. INTENDED USE, SCOPE, AND PRINCIPLE |  |  |  |
| --- | --- | --- | --- |
| 11. | Adequate Number of Tests | Article 20.3(k) | CONTENTS |

The detailed description of content is represented in Tables 1.

| Reagent | Description | Total volume | Amount |
| --- | --- | --- | --- |
| Transport medium | Colorless transparent liquid | 50 mL (500 μL in each) | 100 tubes |

Table 1: STOR F Transport Medium content

All components are ready to use and do not require additional preparation for operation.

The kit in packages S, tubes, and S, screw cap tubes, are intended for single use and designed for 100 tests for STOR F Transport Medium. The kit in package N is designed for 100 tests for STOR F Transport Medium.

REAGENTS AND EQUIPMENT REQUIRED BUT NOT PROVIDED

Specimen Collection

- Sterile single use swabs, cytobrushes, cotton swabs e.t.c for sampling of biomaterial;

- 1.5 mL tubes (only for work with package N).

Preparation For DNA Extraction

- Biological safety cabinet class II;

- Refrigerator;

- Vortex mixer;

- High speed centrifuge (RCF(g) 13000);

- Tube rack for 1.5 mL tubes;

- Single channel pipettes (dispensers covering 100-1000 μL volume range);

- RNase and DNase free filtered pipette tips (volume 1000 μL);

- NA extraction kit;

- Container for used pipette tips, tubes and other consumables;

- Powder-free surgical gloves;

- Disinfectant solution.

| III. SAFETY AND PRECAUTIONS |  |  |  |
| --- | --- | --- | --- |
| 12. | Warnings and Precautions | Article 20.4.1(n) | All necessary general warnings for safe use |
| 13. | Biohazard Warning | Article 20.4.1(p) | Specimens pose a potential infection risk |
| 14. | Hazardous Chemical Substances | Article 20.4.1(o) | If kit components contain hazardous substances |
| 15. | Restriction on Use | Article 20.1(f) | For use only by trained laboratory personnel |

Only personnel trained in the methods of molecular diagnostics and the rules of work in the clinical and diagnostic laboratory are allowed to work with the kit.

Handle and dispose all biological samples, reagents and materials used to carry out the assay as if they were able to transmit infective agents. The samples must be exclusively employed for certain type of analysis. Samples must be handled under a laminar flow hood. Tubes containing different samples must never be opened at the same time. Pipettes used to handle samples must be exclusively employed for this specific purpose. The pipettes must be of the positive dispensation type or be used with aerosol filter tips. The tips employed must be sterile, free from the DNases and RNases, free from DNA and RNA. The reagents must be handled under a laminar flow hood. The reagents required for amplification must be prepared in such a way that they can be used in a single session. Pipettes used to handle reagents must be exclusively employed for this specific purpose. The pipettes must be of the positive dispensation type or be used with aerosol filter tips. The tips employed must be sterile, free from the DNases and RNases, free from DNA and RNA. Avoid direct contact with the biological samples reagents and materials used to carry out the assay. Wear powder-free surgical gloves. Wear protective clothing (work clothes and personal protective equipment) working with microorganisms classified as particularly pathogenic. The protective clothing and personal protective equipment must comply with the work to be performed and health and safety requirements. Avoid producing spills or aerosol. Any material being exposed to biological samples must be treated for at least 30 minutes with disinfecting solution or autoclaved for 1 hour at 121 °C before disposal.

Molecular biology procedures, such as nucleic acids extraction, reverse transcription, PCR-amplification and detection require qualified staff to avoid the risk of erroneous results, especially due to the degradation of nucleic acids contained in the samples or sample contamination by amplification products.

All the liquid solutions are designed for single use and can not be used more than once in amplification reactions. Plastic tubes do not contain phthalates. Do not breathe gas/fumes/vapor/spray produced by the components of the kit. Do not eat/drink components of the kit. Avoid contact with eyes. Only use the reagents provided in the kit and those recommended by manufacturer. Do not mix reagents from different batches. Do not use reagents from third party manufacturers’ kits. All laboratory equipment, including pipettes, test tube racks, laboratory glassware, lab coats, bouffant caps, etc., as well as reagents should be strictly stationary. It is not allowed to move them from one room to another. Equip separate areas for the extraction/preparation of amplification reactions and for the amplification/detection of amplification products. Never introduce an amplification product in the area designed for extraction/preparation of amplification reactions. Wear lab coats, gloves and tools, which are exclusively employed for the extraction/preparation of the amplification reaction and for the amplification/detection of the amplification products. Never transfer lab coats, gloves and tools from the area designed for amplification/detection of the amplification products to the area designed for extraction/preparation of amplification reactions.

Emergency actions

Eye Contact: If any component of this kit enters the eyes, wash eyes gently under potable running water for 15 minutes or longer, making sure that the eyelids are held open. If pain or irritation occurs, obtain medical attention.

Skin Contact: If any component of this kit contacts the skin and causes discomfort, remove any contaminated clothing. Wash affected area with plenty of soap and water. If pain or irritation occurs, obtain medical attention.

Ingestion: If any component of this kit is ingested, wash mouth out with water. If irritation or discomfort occurs, obtain medical attention.

| III. WARNINGS AND PRECAUTIONS |  |  |  |
| --- | --- | --- | --- |
| 16. | Special Storage Conditions | Article 20.4.1(j) | Required temperature range for reagent storage |

Expiry date – 12 months from the date of production.

The STOR F Transport Medium must be transported in thermoboxes with ice packs by all types of roofed transport at temperatures inside the thermoboxes corresponding to storage conditions of the kit components.

It is allowed to transport the kit in thermoboxes with ice packs by all types of roofed transport at temperatures from 2°C to 25°C inside the thermoboxes, but for no longer than 14 days.

Kits transported with violation of temperature conditions must not be used.

All components of the STOR F Transport Medium must be stored in a refrigerator or a cooling chamber at temperatures from 2°C to 8°C over the storage period.

The excessive temperature and light can be detrimental to product performance.

Shelf-life of the kit following the first opening of the primary container: the components of the kit must be stored in a refrigerator or a cooling chamber at temperatures from 2°C to 8°C over the storage period.

The kit stored under undue regime must not be used. An expired STOR F Transport Medium must not be used.

We strongly recommend to follow the given instructions in order to obtain accurate and reliable results.

The conformity of the STOR F Transport Medium to the prescribed technical requirements is subject to compliance of storage, transportation and handling conditions recommended by manufacturer.

| IV. TEST PROCEDURE AND QUALITY CONTROL |  |  |  |
| --- | --- | --- | --- |
| 17. | Sample Type | Article 20.4.1(d) | Type of specimen accepted for testing |

The STOR F Transport Medium is designed to transport and storage the scrapes/smears of epithelial cells from urogenital tract, oropharynx, nasopharynx, rectum, skin, conjunctiva of the eye.

General requirements

Remove free separable mucus with a sterile cotton prior to sampling. In case of sampling from several locations, repeat the procedure several times, each time taking a new swab into new different tube.

To prevent contamination, open the tube, add sample, then close the tube before proceeding to the next sample.

STOR F Transport Medium is ready for use. Mark necessary number of 1.5 mL tubes. Add 500 μL of transport medium into marked tubes. Close the tubes tightly.

Prepared tubes should be stored at temperatures from 2°C to 8°C for the entire shelf life, and used when necessary.

Scrapes/smears of epithelial cells sampling

Order of taking:

1. Open the 1.5 mL tube with STOR F Transport Medium.

2. Scrape epithelial cells from the corresponding biotope (i.e. urogenital tract, oropharynx, nasopharynx, rectum, skin, conjunctiva of the eye) with a sterile swab.

3. Put the swab into the tube with transport medium and rinse it thoroughly. Avoid spraying of solution.

4. Remove swab from solution, press it to the wall of tube and squeeze the rest of the liquid. Throw out the swab.

5. Close the tube tightly and mark it.

Transportation and storage of the samples

Transport and store samples in STOR F Transport Medium at temperatures from 2°С to 8°С for no longer than 7 days.

Transport and store samples in STOR F Transport Medium at temperatures from 18°С to 25 °С is acceptable for no longer than 48 hours.

| IV. TEST PROCEDURE AND QUALITY CONTROL |  |  |  |
| --- | --- | --- | --- |
| 18. | Refer to Instructions for Use | Article 20.3(j) | Must refer to instructions for detailed procedure |
| 19. | Procedure Steps | Article 20.4.1(l) | Flowchart from nucleic acid extraction to qPCR cycling |

General Recommendations

1. Use only disposable RNase and DNase free filtered pipette tips.

2. When adding the solution into sample, do not touch the walls of the tubes with the tip. If the tip has touched the wall of the tube, change the tip. Tip should be changed each time when you take out solution from the tube with sample.

3. To prevent contamination, open the tube, add sample/reagent, then close the tube before proceeding to the next sample/reagent.

Preparation for NA Extraction

1. Centrifuge samples in transport medium at RCF(g) 13,000 for 10 minutes.

2. Remove supernatant, leaving the volume recommended by NA extraction kit (precipitate+liquid fraction) from this kind of sample.

ATTENTION! If for sample pretreatment centrifugation is not needed, p.8.2.1 and 8.2.2 are not performed.

ATTENTION! For scrape/smear from oropharynx and nasopharynx for subsequent analysis of viral RNA, preliminary centrifugation is not required.

3. Perform NA extraction according to NA extraction kit user manual.

4. To  make  a  negative  control  perform  p.  8.2.1  and  8.2.2  for  tube  with STOR F Transport Medium which does not contain sample. Or use necessary volume of saline as sample according to NA extraction kit manual.

STOR F Transport Medium is compatible with any NA extraction methods.

| IV. TEST PROCEDURE AND QUALITY CONTROL |  |  |  |
| --- | --- | --- | --- |
| 20. | Quality Control (QC) Procedure | Article 20.4.1(m) | Details for use of Positive, Negative and Internal Control (IC) |

| V. RESULTS AND PERFORMANCE |  |  |  |
| --- | --- | --- | --- |
| 21. | Interpretation of Results | Article 20.4.1(m) | Ct cut-off values and result table for each pathogen |

| V. RESULTS AND PERFORMANCE |  |  |  |
| --- | --- | --- | --- |
| 22. | Analytical Performance Data | Article 20.4.1(q) | Technical parameters such as LOD, Specificity, Precision, Linear Range. |
| 23. | Clinical Performance Data | Article 20.4.1(q) | Diagnostic Sensitivity, Specificity, and expected values (prevalence) in pediatric population. |

| Parameter | Storage conditions |  |
| --- | --- | --- |
|  | 48 hours under temperatures from 18°С to 25°С | 7 days under temperatures from 2°С to 8°С |
|  | Average, 95% CI* |  |  |
| Preservation of human genomic DNA (%),calculated for 50 samples | 87.9 (83.9 – 92.0) | 86.3 (80.2 – 92.5) |  |
| Preservation of viral nucleic acids (detected by AVRI kit11) (%), calculated for 21 samples | 97.6 (91.5 – 100) | 79.9 (67.8 – 92.1) |  |
| *95% CI - 95% confidence interval |  |  |  |

The preservation of nucleic acids in biomaterial samples in the STOR F Transport Medium was estimated as the relative amount of nucleic acids (%) in the samples after storage in the transport medium at different temperature-time regimes compared to samples without storage, calculated using the indicator cycle comparison method (ΔCp) in real-time PCR reaction.

| V. RESULTS AND PERFORMANS |  |  |  |
| --- | --- | --- | --- |
| 24. | Retesting Requirements | Article 20.4.1(m) | How to repeat invalid or equivocal results. |

| VI. ADDITIONAL INFORMATION |  |  |  |
| --- | --- | --- | --- |
| 25. | Disposal of Used Materials | Article 20.4.1(t) | Instructions for medical and hazardous waste management |

Remove waste materials (tubes, tips) only in a special closed container containing a disinfectant solution. Work surfaces, as well as rooms where NA extraction and PCR are performed, must be irradiated with bactericidal irradiators for 30 minutes before and after the work. Waste materials are disposed of in accordance with local and national standards. All surfaces in the laboratory (work tables, test tube racks, equipment, etc.) must be treated daily with disinfecting solution.

| VI. ADDITIONAL INFORMATION |  |  |  |
| --- | --- | --- | --- |
| 26. | Revision Date of Instructions for Use | Article 20.4.1(x) | 2025-11-25 |
| 27. | Standards with which Compliance is Claimed | Article 20.4.1(v) | ISO 15223-1, ISO 18113 Reference to relevant standards such |

| VI. ADDITIONAL INFORMATION |  |  |  |
| --- | --- | --- | --- |
| 28. | Check Before Use | Article 20.4.1(r) | Actions in case of damaged packaging or reagent deterioration |

Do not use the kit:

- When the transportation and storage conditions are breached;

- When the reagents’ appearance does not respond to the kit passport;

- When the kit components packaging is breached;

- After the expiry date provided.

Significant health effects are NOT anticipated from routine use of this kit when adhering to the instructions listed in the current manual.

| VI. ADDITIONAL INFORMATION |  |  |  |
| --- | --- | --- | --- |
| 29. | Technical Support Contact | - | Contact details for questions and technical issues. |

For technical support, please contact our dedicated technical support team by email: support@genomer.com.tr or Phone: +90 530 462 2637.

VII. SYMBOLS/PICTOGRAMS

|  | In vitro diagnostic medical device |  | Manufacturer |
| --- | --- | --- | --- |
|  | Temperature limit |  | Date of manufacture |
|  | Contains sufficient for <n> tests |  | Consult instructions for use |
|  | Use-by date |  | Catalogue number |
|  | Batch code |  | Version |
|  | Do not re-use |  | Non-sterile |

# REVISION HISTORY

| Revision No. | Rev. Date | Description |
| --- | --- | --- |
| 00 | 14.11.2025 | First Issue |

# DOCUMENT APPROVAL

| Prepared By | Approved By |
| --- | --- |
| Quality Management Representative | General Manager |
| Elife ÖZDEMİR | Oktay ÖZTÜRK |